## Version
Release Date | Effective Date
Sep 18 2026 | Sep 18 2026
General Requirements:
1 | Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees.
2 | Provision of regular updates and reports about the product to DOH upon request.
3 | Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector.
4 | Compliance with the relevant laws and standards of the Department of Health.
5 | Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
6 | Please note that this product will be monitored and followed up by the Department of Health. In the event that any complications or side effects are proven as a result of using this product, the Department of Health has the authority to withdraw, suspend and investigate this product in the health sector of the Emirate of Abu Dhabi.
7 | Review Coverage with Healthcare Facility if applicable

## Approved HTA Medical Products
Decision Date | Reference No. | Technology Title, Model/version no. | Agent / Hospital | MANUFACTURER | Country of origin | Tech Type | Tech Category | Condition of Use
2026-09-09 | HTA26073 | Magstim TMS Therapy — Rapid² System | BAYAN MEDICAL COMPANY LLC | Magstim Company Limited | United Kingdom | Medical Device | Therapeutic | Approved for adoption for the treatment of Major Depressive Disorder in adult patients who have failed to achieve satisfactory improvement from prior antidepressant medication in the current episode. Recommendations: 1.The Magstim Rapid² Therapy System is recommended for use in licensed hospitals or outpatient psychiatric/neuromodulation clinics for repetitive transcranial magnetic stimulation of adult patients with treatment-resistant Major Depressive Disorder where clinically indicated. 2.Use must be restricted to licensed physicians (psychiatrists) trained in TMS, with delivery by trained technicians or nurses under documented physician supervision. 3.The system must be used only within its FDA-cleared indication and parameter set: 10 Hz, 120% motor threshold, 3,000 pulses per session, ~37.5 minutes, 5 sessions weekly for up to 6 weeks, targeting the frontal cortex. 4.Patient selection must be documented by a qualified psychiatrist, confirming failure of at least one adequate antidepressant trial in the current episode. 5.A pre-treatment screening questionnaire must be completed for every patient, with all Operating Manual contraindications applied as hard stops in the electronic order set. 6.Motor threshold must be determined and documented by a competent operator at the start of each treatment course. 7.Hearing protection must be worn by the patient for cortical stimulation above 70% power, and stimulation dose limits must not be exceeded (≤250 stimuli/minute, ≤4,000/hour, ≤24,000 per 24 hours at 100% power). 8.The facility must maintain a biomedical engineering electrical safety and preventive maintenance programme; no modification of the equipment is permitted and servicing is restricted to qualified personnel. 9.Seizures, hypomania, and all serious adverse events or device malfunctions must be reported through post-market surveillance and DOH pharmacovigilance channels. Clinical Use Overview •Population / Intended User — Adults aged 18 years and over with Major Depressive Disorder who have failed to achieve satisfactory improvement from prior antidepressant medication in the current episode. •To be Performed By — Licensed physician (psychiatrist) trained in TMS; delivery by a trained TMS technician or nurse under documented physician supervision. •Clinical Setting — Licensed hospital, outpatient psychiatric clinic, or neuromodulation clinic with on-site seizure management capability. •Condition of Use — Prescription use only; cleared stimulation parameters (10 Hz, 120% motor threshold, 3,000 pulses per session, 5 sessions weekly for up to 6 weeks); mandatory pre-treatment screening questionnaire; hearing protection above 70% power; dedicated earthed mains supply with supplied filtered leads and earth straps. Contraindications: •Used on or in the vicinity of patients or users with cardiac demand pacemakers, implanted medication pumps, cochlear devices, implanted defibrillators and/or implanted neurostimulators. •Used on or in the vicinity of patients with implanted metal objects. •Used on patients where the skin in the area to be contacted is broken. •Used on patients who suffer from multiple sclerosis. •Used on those with large ischemic scars. •Used on pregnant women. •Used on infants under the age of 2 years. •Used on epileptic individuals. •Used on those with a family history of convulsions. •Used on individuals with brain lesions that could affect seizure threshold. •Used on individuals taking tricyclic antidepressants, neuroleptic agents or any other drug that could lower seizure threshold. •Used on individuals suffering from sleep deprivation during rTMS procedures. •Used on individuals with a heavy consumption of alcohol or those using epileptogenic drugs. •Used on individuals with severe heart disease or with increased intracranial pressure, or on those who have uncontrolled migraines. •Do not use if there is any sign of external damage to the system, or if any parts are damp or wet. •Do not use in an explosive atmosphere, around explosive gases, or in the presence of flammable anaesthetics. •Do not use for cortical stimulation when stimulating at 100 Hz — peripheral use only at that frequency. •Do not use outside the FDA-cleared indication or cleared parameter set. •Do not use by untrained personnel or without documented physician supervision. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26062 | STANDARD M10 REAL TIME GENE AMPLIFIER | Babirus Medical Equipment Trading LLC | SD BIOSENSOR, Inc. | Republic of Korea | Medical Device | Diagnostics | Approved - The STANDARD M10 is an in vitro diagnostic device that measures quantitative or qualitative biomarkers of human specimens such as blood, urine, and nasal mucus in a laboratory or POCT environment. STANDARD M10 consists of STANDARD M10 Console and STANDARD M10 Module (Cartridge analysis function). The STANDARD M10 Console is a device that operates the STANDARD M10 Module based on the information obtained by scanning the 2D barcode of the cartridge. The M10 Console displays the results of monitoring and diagnosis of the inserted specimen. It must be used with designated test cartridges produced by SD Biosensor Co., Inc. For details on specific tests, refer to the instructions included in each of test cartridge packages. Recommendation: -Adopt the system for appropriate molecular diagnostic testing: The STANDARD M10 may be implemented as an automated real-time PCR platform for molecular diagnostic testing within its intended use. The specific clinical application is determined by the assay cartridge used and its respective authorized intended purpose. -Use only appropriate and authorized assay cartridges: The STANDARD M10 should be used only with STANDARD M10 assay applications intended for use with the system. Each assay cartridge should be considered according to its own intended use, specimen type, performance characteristics, limitations, and applicable regulatory authorization. The system automatically recognizes the assay to be performed based on the cartridge used. -Ensure laboratory verification and personnel training: Before routine implementation, the healthcare facility should verify system performance under local operating conditions and ensure that personnel are appropriately trained in system operation, cartridge handling, specimen processing, result interpretation, quality control, and troubleshooting. -Maintain quality assurance and regulatory compliance. -Ensure appropriate calibration and maintenance. Clinical overview: Population patient population is determined by the intended use of the individual assay cartridge. To be performed by Qualified and appropriately trained laboratory or healthcare professionals, according to the intended use and requirements of the specific assay cartridge. Clinical setting Laboratory or clinical diagnostic settings, according to the authorized intended use of the specific assay cartridge. Condition of use The STANDARD M10 is an automated real-time PCR molecular diagnostic system intended for use with compatible STANDARD M10 assay applications. Each assay cartridge has its own intended purpose, specimen requirements, clinical performance, limitations, and result-interpretation requirements; therefore, approval of the analyzer should not be interpreted as approval of every individual assay or clinical indication. AAMEN Approved: -Health data and PII shall not be transferred or made accessible outside the UAE. -Support (L1, L2 and L3) shall have no visibility into PHI/PII. -L1 and L2 support shall be within UAE. -System logs must be made available for the health entity. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26059 | CMG – Compressomyography for Neuromuscular Blockade Monitoring - Vision DUO | Emirates & World Medical Supplies | RGB Medical Devices | Spain | Medical Device | Patient Monitoring During Anesthesia | Conditionally Approved for adoption for adult patients and to monitor neuromuscular transmission and non invasive blood pressure of adult patients and issue alarms related with these physiological parameters. It is not a therapeutic device. Recommendations 1. Vision DUO is recommended for hospital use for quantitative neuromuscular transmission and non-invasive blood-pressure monitoring in adult patients. 2. Use shall be restricted to trained and competent healthcare professionals, principally anesthesiologists and appropriately trained anesthesia/perioperative nursing staff. 3. The device shall not be used in pediatric patients. 4. Vision DUO shall be used as a clinical monitoring tool and shall not independently determine administration of neuromuscular blocking agents, reversal agents or extubation. 5. Facilities shall incorporate quantitative neuromuscular monitoring into perioperative protocols whenever neuromuscular blocking agents are administered, consistent with contemporary anesthesia guidelines using Vision DUO or another validated quantitative monitoring method. 6. Healthcare professionals must recognize that Vision DUO evaluates neuromuscular response at an upper-arm muscle group and that its TOF-ratio values are not directly interchangeable with adductor-pollicis EMG/AMG values. 7. Particular caution shall be applied where accurate confirmation of complete neuromuscular recovery is clinically critical. Use of an appropriate adductor-pollicis quantitative monitoring method should be considered when clinically required. 8. Correct cuff size shall be selected according to measured limb circumference before monitoring. 9. The cuff shall not be placed on a limb: • With an arterial catheter. • With an intravenous infusion. • With compromised circulation. • With injured, inflamed or damaged skin. • On the side of previous mastectomy or lymph-node clearance where avoidable. 10. The patient and cuff site shall be periodically evaluated during repeated monitoring because frequent cuff inflation or nerve stimulation may contribute to ischemia, purpura, neuropathy or impaired circulation. 11. Patients shall be anesthetized before electrical NMT stimulation is initiated because nerve stimulation can be painful in conscious patients. 12. NMT results should be interpreted cautiously in patients with nerve damage, myasthenia gravis, Bell's palsy or other neuromuscular disorders. 13. The device has not been validated in the obese population. 14. The monitor shall not be used during MRI examinations or within the MR environment. 15. Only RGB Medical Devices-approved cuffs, cables and accessories shall be used. 16. The device and its accessories shall be inspected before use, including the cuff, cables, electrical connections and monitor condition. 17. Staff shall receive initial and periodic training covering: • Correct cuff selection and positioning. • NMT-Cuff / compress myography principles. • TOF, ST and PTC interpretation. • AUTO-PILOT operation. • Differences between upper-arm CMG and adductor-pollicis monitoring. • Recognition of residual neuromuscular blockade. • Device alarms and troubleshooting. • NIBP measurement limitations. 18. Real world evidence shall be collected following adoption to address residual uncertainties related to clinical performance, recovery interpretation and patient outcomes when Vision DUO is used in routine practice. RWE collection should focus on recovery phase interpretation, use of complementary monitoring, residual neuromuscular blockade in PACU, and perioperative respiratory outcomes, particularly in patient subgroups under represented in pre market studies (Obese patients). Specific clinical outcomes and quality indicators to be monitored following adoption, include: • Percentage of patients receiving quantitative NMT monitoring when NMBAs are administered. • TOF status prior to extubation. • Requirement for complementary monitoring. • Residual neuromuscular blockade in PACU. • Respiratory events/reintubation. • Device-related adverse events. • Cuff-related skin or circulation complications. • Device malfunction and measurement failure. 20. Device malfunctions, serious incidents and medical-device vigilance events shall be reported in accordance with DoH requirements. Clinical Use Overview Population / Intended User: Adult patients requiring monitoring of neuromuscular transmission and/or non-invasive blood pressure, particularly adults receiving neuromuscular blocking agents during general anesthesia. The current IFU does not authorize pediatric monitoring. To be Performed By • Anesthesiologists. • Other responsible physicians appropriately trained in NMT monitoring. • Trained nurses working within the relevant hospital unit and under appropriate clinical governance. • Other trained medical professionals authorized by the healthcare facility. The device must never be directly operated by the patient. Clinical Setting • Operating rooms. • Post-anesthesia/recovery units where clinically appropriate. • Intensive care units. • Resuscitation/critical-care areas. • Other appropriate licensed hospital areas where NMT or NIBP monitoring is clinically required. Condition of Use 1. Adult patient. 2. Use within a licensed healthcare institution. 3. Qualified and trained professional responsible for monitoring. 4. Appropriate cuff size selected according to limb circumference. 5. Cuff correctly positioned and electrodes aligned with the intended nerve. 6. Patient anesthetized before NMT electrical stimulation. 7. Limb free from contraindicating vascular access, damaged skin or compromised circulation. 8. TOF/ST/PTC parameters interpreted according to the manufacturer's IFU. 9. Recognition that upper-arm Vision DUO recovery occurs earlier than adductor-pollicis recovery in some patients. 10. Complementary clinical assessment performed before extubation. 11. Alternative/reference quantitative monitoring available when measurements are uncertain or inconsistent with the patient's clinical status. 12. Device alarms appropriately configured and audible. 13. Manufacturer-approved accessories only. 14. Routine cleaning, preventive maintenance and electrical-safety testing undertaken according to the IFU. 15. Device used outside the MRI environment. 16. Medical-device incident reporting and quality-monitoring mechanisms established. Contraindications / Limitations 1. Pediatric patients,. 2. Direct patient/self-use. 3. Home use. 4. Use in an MRI environment. 5. Use in the presence of flammable anesthetic mixtures with air, oxygen or nitrous oxide. 6. Nerve stimulation in a conscious/non-anesthetized patient. 7. Placement of stimulation electrodes on the head, neck or thorax. 8. Cuff placement on a limb with an arterial catheter. 9. Cuff placement on a limb with an intravenous infusion. 10. Cuff placement on a limb with compromised circulation. 11. Cuff placement over injured, inflamed or significantly damaged skin. 12. Ipsilateral cuff placement following mastectomy or lymph-node clearance should be avoided. 13. Use of an incorrect cuff size. 14. The device has not been validated in the obese population. 15. Measurements may be unreliable during significant tremor or convulsions. 16. NMT measurements may be inaccurate in nerve damage, myasthenia gravis, Bell's palsy or other neuromuscular disorders. 17. TOF-ratio assessment has limitations with depolarizing neuromuscular blocking agents. 18. Vision DUO TOFR values must not be assumed to be directly interchangeable with adductor-pollicis AMG, EMG or mechanomyography values. 19. A Vision DUO TOFR ≥0.9 must not, by itself, be considered proof that residual neuromuscular blockade has been completely excluded. 20. If the displayed NMT or NIBP measurement is inconsistent with the patient's clinical condition, it should be verified using another clinically accepted measurement method. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26058 | Allplex™ HPV28 Detection | Seegene Medical Equipment Trading LLC | Seegene Inc. | South Korea | Diagnostic Test | Diagnostics | Approved for Allplex™ HPV28 Detection is in vitro diagnostic medical device designed for qualitative detection of human papillomaviruses in cervical specimens self-collected vaginal specimens. Allplex™ HPV28 Detection consists of two PCR reactions (A and B MOM). A MOM is a multiplex assay that permits the simultaneous amplification of the target DNA of 14 high-risk human papillomaviruses. B MOM is a multiplex assay that permits the simultaneous amplification of the target DNA of 5 high-risk and 9 low-risk human papillomaviruses. A MOM 14 high-risk HPV types Types (16, 18, 31, 33, 35, 39, 45, 51, 52, 56, 58, 59, 66, 68) B MOM 9 low-risk HPV types (6, 11, 40, 42, 43, 44, 54, 61, 70) 5 high-risk HPV types (26, 53, 69, 73, 82) Recommendations: • To be performed by qualified and trained laboratory professionals only. • The assay should be integrated into the established cervical cancer screening pathway and interpreted in conjunction with the patient’s clinical history, previous screening results, cervical cytology and other relevant clinical findings. • Only validated cervical specimens and self collected vaginal specimens should be tested; use with other specimen types is not recommended unless separately validated. • The test should not be used as a standalone method for the diagnosis of cervical precancer or cancer. Positive high risk HPV results should be managed through the appropriate follow up pathway, including cervical cytology, repeat HPV testing, colposcopy and/or biopsy, as clinically indicated. • Negative results do not completely exclude HPV infection, particularly in cases of inadequate specimens, viral loads below the limit of detection, or inappropriate specimen collection, transport or storage conditions. • Laboratories should verify compatibility with manufacturer specified nucleic acid extraction systems • Strict procedures must be maintained to prevent sample and amplicon contamination, including the use of separated work areas, filter tips, appropriate positive and negative controls, and validated quality control procedures. • Storage conditions, specimen handling requirements, reagent expiry, infection control measures and freeze–thaw limitations must be strictly followed in accordance with the manufacturer’s instructions. • Clinical performance and post market outcomes, including invalid results, genotype distribution, concordance with follow up cytology or histology, discordant findings, repeat testing rates, complaints and adverse incidents, should be routinely monitored and documented. Clinical Use Overview: Population / Intended User: Women aged 25–65 years undergoing HPV testing or cervical cancer screening. To be Performed By: Qualified laboratory and healthcare professionals. Clinical Setting: Hospitals / Laboratories Condition of Use: Allplex™ HPV28 Detection is a multiplex real time PCR assay intended for the qualitative detection and individual genotyping of 28 HPV types, including 19 high risk and nine low risk genotypes, in cervical and self collected vaginal specimens. The assay should be used within an established HPV screening or diagnostic pathway and not as a standalone test for cervical cancer. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26057 | Khan Kinetic Treatment Device (KKT-M2) | KKT Orthopedic Spine Center | Optima Health Solutions International Corp. | Canada | Medical Device | Therapeutic | Conditional Approval for use as an adjunctive conservative therapy to be used in the aid of management of chronic pain due to non-congenital defects. The device can be used as part of a series of steps in the total care of the patient. The procedure involves the use of diagnostic imaging that qualifies the misalignment between vertebrae. The treatment is then administered using the KKT M2 device to deliver precise impulses at a required vector configuration. Recommendations 1.KKT-M2 is recommended as an adjunctive/second-line conservative treatment for appropriately selected adults with chronic mechanical neck or low back pain who remain symptomatic despite appropriate first-line conservative management. 2.KKT shall not replace established first-line management including patient education, therapeutic exercise, physiotherapy and other guideline-recommended conservative treatment. 3.Use shall be restricted to licensed healthcare professionals who have completed appropriate KKT training and demonstrated competency in device operation and spinal assessment. 4.Patients shall undergo comprehensive clinical assessment before treatment, including medical history, musculoskeletal examination, neurological assessment and appropriate imaging where clinically indicated. 5.Treatment shall be restricted to patients without red-flag conditions requiring urgent diagnostic, neurological or surgical evaluation. 6.Treatment shall not be performed in patients with spinal fracture, spinal infection, malignancy involving the spine, destructive spinal lesions, severe instability or progressive neurological emergencies. 7.Treatment planning and delivery shall be individualized according to clinical and radiographic assessment and undertaken within established device parameters. 8.The treatment course should be periodically reassessed and continued only where there is demonstrable clinical benefit. 9.Clinical outcomes should be assessed using validated measures such as: o Pain intensity score. o Neck Disability Index for cervical pain. o Oswestry Disability Index for low back pain. o Functional status and activity. o Analgesic utilization. 10.Clinical outcomes should be documented at baseline, following the initial treatment course and during subsequent follow-up where treatment is continued. 11.Adverse events, treatment discontinuation, lack of clinical response, escalation to spinal injections or surgery and device-related incidents should be monitored and documented. 12.Claims of intervertebral disc regeneration, reversal of degenerative disc disease or structural disease modification should not be made, as the current biological evidence is preclinical. 13.The technology should not be promoted as superior to established physiotherapy, exercise-based care or multidisciplinary rehabilitation because comparative evidence is insufficient to establish superiority. 14.KKT adoption should include periodic clinical audit and review of real-world outcomes for UAE population, particularly pain, functional improvement, number of treatment sessions, durability of benefit, analgesic use and subsequent escalation to invasive interventions. 15.Device malfunctions and serious incidents shall be reported in accordance with applicable DoH and UAE regulatory requirements. 16.Any material change to the device, intended use, treatment protocol or treatment-delivery software should undergo appropriate regulatory and HTA reassessment. Clinical Use Overview Population / Intended User: Adults with chronic mechanical neck pain or chronic low back pain, particularly those with persistent symptoms despite appropriate first-line conservative management and who are considered suitable following clinical and radiographic assessment. To be Performed By: Licensed healthcare professionals appropriately trained and competent in spinal assessment, patient selection and KKT-M2 operation. Clinical Setting: Licensed outpatient musculoskeletal, physical medicine, rehabilitation, orthopedic, spine or pain-management healthcare facility with appropriate clinical governance and referral pathways. KKT Orthopedic Spine Centre Abu Dhabi has provided an undertaking confirming intended clinical use following DoH approval and commitment to trained professional operation and incident reporting. Condition of Use: 1.Adult patient with chronic mechanical neck or low back pain. 2.Appropriate clinical assessment completed. 3.Red-flag spinal pathology excluded. 4.Neurological assessment completed. 5.Radiographic assessment performed where required for KKT treatment planning. 6.Persistent symptoms despite appropriate conservative management. 7.Treatment delivered by a trained and licensed healthcare professional. 8.Individualized treatment parameters based on clinical/radiographic assessment. 9.Clinical response reassessed after the initial treatment course. 10.Continued treatment only where measurable clinical benefit is demonstrated. 11.Appropriate referral pathway available for patients requiring specialist, interventional or surgical management. Contraindications / Limitations: 1.Acute spinal fracture. 2.Spinal infection. 3.Malignancy involving the spine. 4.Destructive spinal lesion. 5.Severe or clinically significant spinal instability. 6.Progressive neurological deficit or neurological emergency. 7.Congenital spinal conditions outside the authorized indication. 8.Acute condition requiring surgical or urgent specialist management. 9.Patients in whom mechanical spinal treatment is considered clinically unsafe. 10.Use as a replacement for appropriate diagnostic investigation. 11.Use as a substitute for first-line evidence-based conservative management. 12.Use to claim or treat human intervertebral disc regeneration on the basis of current preclinical evidence alone. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26054 | m-PIMA™ Analyser | Al Shomoukh Trading for Technical & Medical Supplies | ABBOTT RAPID DIAGNOSTICS JENA GMBH,GERMANY | Germany | Diagnostic Test | Diagnostics | Approved for : The m-PIMA™ Analyzer is a portable automated bench-top analyzer for processing m-PIMA™ test cartridges. The m-PIMA™ Analyzer is suited to be used in laboratory as well as in non-laboratory environments. The instrument is classified as in vitro diagnostic device. The following assays used with the platform: · m-PIMA™ HIV-1/2 Detect – qualitative molecular detection of HIV-1 and HIV-2. The m-PIMA'" HIV-1/2 Detect test is intended to be used as an aid in the diagnosis of HIV infection in pediatric and adult individuals, or as a confirmatory assay when samples have already been tested using alternative methods (e.g. serological assays to screen for evidence of HIV infection). · m-PIMA™ HIV-1/2 Viral Load – quantitative measurement of HIV-1 and HIV-2 RNA. This test is intended for use in conjunction with clinical presentation and other laboratory markers of disease progress for the clinical management of HIV-I and HIV-2 infected patients. This test can be used to assess patient prognosis by measuring the baseline HIV-I and HIV-2 RNA level or to monitor the effects of antiretroviral therapy by measuring changes in EDTA plasma HIV-I and HIV-2 RNA levels during the course of antiretroviral treatment. Recommendations: 1. The m-PIMA Analyzer is recommended for use in laboratories, HIV clinics and eligible decentralized or near-patient settings where rapid molecular HIV testing is clinically required. 2. The analyzer should be operated only by trained and competent healthcare, laboratory or appropriately trained non-laboratory personnel. 3. Each cartridge should be used only for its approved intended purpose. 4.The m-PIMA HIV-1/2 Detect test is not intended to be used as a donor screening test for HIV. 5. The m-PIMA HIV-1/2 Viral Load test is not intended for use as a screening test for the presence of HIV-1 and HIV-2 in blood or blood products or as a diagnostic test to confirm the presence of HIV-1 and HIV-2 infection. 6. Viral-load results should be interpreted together with the patient’s clinical presentation, treatment history and other laboratory markers. 7. Invalid, unexpected or discordant results should be repeated or referred to a centralized laboratory. 8. Appropriate internal quality control, external quality assessment, operator-competency assessment and invalid-result monitoring should be established. 9. Only approved specimen types, collection methods, sample volumes and storage conditions should be used. 10. Infection-control procedures, personal protective equipment and biohazard-waste requirements must be followed. 11. The analyzer should be operated within the specified environmental conditions and connected to external power whenever possible. 13. Post-adoption monitoring should include turnaround time, invalid-test rate, repeat-test rate, cartridge wastage, analyser downtime, result concordance and clinical incidents. Clinical Use Overview: Population / Intended User: The m-PIMA'" HIV-1/2 Detect test is intended to be used as an aid in the diagnosis of HIV infection in pediatric and adult individuals, or as a confirmatory assay when samples have already been tested using alternative methods (e.g. serological assays to screen for evidence of HIV infection). The m-PIMA HIV -1/2 Viral Load This test can be used to assess patient prognosis by measuring the baseline HIV-I and HIV-2 RNA level or to monitor the effects of antiretroviral therapy by measuring changes in EDTA plasma HIV-I and HIV-2 RNA levels during the course of antiretroviral treatment. To be Performed By: Trained healthcare professionals, laboratory personnel or appropriately trained healthcare workers. Clinical Setting: Laboratories, HIV clinics, hospitals, decentralized healthcare centers and eligible near-patient testing locations. Condition of Use: Professional in vitro diagnostic use with compatible single-use m-PIMA cartridges, within the environmental and operating conditions specified by the manufacturer. Contraindications / Limitations: · Not intended for self-testing. · The viral-load cartridge is not intended for blood screening. · Results should not be interpreted without relevant clinical and laboratory information. · Damaged, expired, wet or incorrectly stored cartridges must not be used. · Invalid or discordant results require repeat or referral testing. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26053 | IRON‑qPCR™ (Fully Automated POCT qPCR System (A‑2600) | Al Amin Medical Instrument Company | Bioneer Corporation | Republic of Korea | Medical Device | Diagnostics | Approved – IRON-qPCR™ is a POCT qPCR device for in vitro diagnostics. It combines a unique format Prep module, sonicator, thermal block and innovative fluorescence detector. After sample injection, if only the kit is installed on the instrument and the program is run, the amount of fluorescence is automatically measured through a fluorescence detector after DNA or RNA extraction. The Prep module extracts DNA or RNA from each solution in the cartridge through the optimal movement path, and at this time, it uses an ultrasonic vibrator to help extract it quickly. The detector monitors in real time the fluorescence emitted as an indicator for the amplified nucleic acid product during each PCR cycle. It can be used for qualitative or quantitative analysis of specific DNA/RNA in a sample. This instrument was developed for users who have experience or have basic knowledge to use real time PCR instrument. Recommendation: 1. Adopt the IRON-qPCR™ A-2600 as a fully automated real-time PCR platform for molecular diagnostic testing within its intended use. 2. Use only compatible BIONEER diagnostic kits intended for use with the IRON-qPCR™ system. 3. Assess each diagnostic kit separately according to its intended use, target analytes, specimen type, patient population, performance, limitations and regulatory authorization. 4. Do not interpret approval of the analyzer as approval of all available assays or clinical indications. 5. Ensure appropriate user training and competency in real-time PCR operation, cartridge handling, specimen processing, result interpretation, quality control and troubleshooting. 6. Verify system performance before routine clinical implementation under local operating conditions. 7. Operate the system according to the manufacturer’s requirements, including an operating temperature of 10–35°C and relative humidity of 20–80% (non-condensing). 8. Maintain appropriate quality assurance, maintenance and incident-reporting procedures in accordance with applicable DoH and laboratory requirements. 9. Use the specific diagnostic kit IFU to determine the clinical indication, specimen type, patient population and interpretation of results for each test performed on the platform. Clinical Overview: Population patient population is determined by the intended use of the individual diagnostic kit used with the IRON-qPCR™ system. To be performed by Qualified and appropriately trained laboratory or healthcare professionals. Clinical setting Laboratory or clinical diagnostic settings, according to the authorized intended use of the specific assay cartridge. Condition of use The IRON-qPCR™ A-2600 is a fully automated real-time PCR molecular diagnostic platform that performs nucleic acid extraction, amplification, fluorescence detection and result reporting using compatible BIONEER diagnostic kits. The system automatically recognizes the protocol associated with the kit through its barcode. Each diagnostic kit has its own intended purpose, target analytes, specimen requirements, patient population, clinical performance, limitations and result-interpretation requirements. Therefore, approval of the IRON-qPCR™ A-2600 platform should not be interpreted as approval of all individual assays or clinical indications available for use with the system. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26052 | Moleculight device | Al Zahrawi Medical Services | MolecuLight Inc. | Canada | Medical Device | Diagnostics | Approved as a handheld imaging tool that allows clinicians diagnosing and treating skin wounds, at the point of care, to: (i) View and digitally record images of a wound, (ii) Measure and digitally record the size of a wound, and (iii) View and digitally record images of fluorescence emitted from a wound when exposed to an excitation light. The fluorescence image, when used in combination with clinical signs and symptoms, has been shown to increase the likelihood that clinicians can identify wounds containing bacterial loads >104 CFU per gram as compared to examination of clinical signs and symptoms alone. The MolecuLightDX device should not be used to rule-out the presence of bacteria in a wound. The MolecuLightDX does not diagnose or treat skin wounds. Recommendations 1. The MolecuLightDX+ is recommended for use in hospitals, wound care clinics, outpatient clinics, and other healthcare facilities involved in the assessment and management of acute and chronic wounds. 2. The device should be operated by trained healthcare professionals in accordance with the manufacturer’s Instructions for Use. 3. The MolecuLightDX+ should be used as an adjunct to routine clinical assessment and should not replace clinical judgment or microbiological culture when clinically indicated. 4. Fluorescence imaging findings should be integrated into wound assessment, debridement, wound sampling, treatment planning, and monitoring of treatment response. 5. Healthcare professionals should receive appropriate training in fluorescence image acquisition and interpretation before clinical use. Clinical Use Overview Population / Intended User * Patients with acute and chronic wounds requiring assessment for elevated bacterial burden. To be Performed By * Trained healthcare professionals involved in wound assessment and management. Clinical Setting * Hospitals. Condition of Use * Non-invasive, point-of-care imaging during routine wound assessment. * Used in accordance with the manufacturer’s Instructions for Use. AAMEN APPROVED: -Health data must not be transferred or made accessible outside the UAE. -Support (L1, L2 and L3) shall have no visibility of PII. -L1 and L2 support shall be within UAE General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26051 | MOSPIRE Human Papillomavirus Self Sample Collection Kit | Omics One Sciences Medical Supplies Trading LLC | Jiangsu Mole Bioscience Co., Ltd. (EU Authorised Representative: SUNGO Europe B.V., Netherlands) | China | Medical Product/ Kit | Diagnostics | Approved for adoption for self-collection, transport and preservation of a cervico-vaginal specimen for laboratory human papillomavirus analysis within an organized cervical cancer screening programme. Recommendations: 1.The MOSPIRE Human Papillomavirus Self Sample Collection Kit is recommended for use in DOH-licensed healthcare facilities and supervised community screening programmes for self-collection of cervico-vaginal specimens for laboratory hrHPV testing, where clinically indicated. 2.Use must be restricted in which the receiving laboratory is DOH-licensed and the hrHPV assay has been validated for self-collected specimens in this preservation medium. Validation evidence must be documented and held on file before first clinical use. 3.Facilities must have a defined, auditable pathway routing every hrHPV-positive woman to clinician-taken triage and colposcopy, with documented communication of the result to the woman. 4.Case selection must be documented at the point of kit issue, confirming the woman is eligible for screening and is not gynecologically unwell, menstruating or otherwise bleeding, pregnant, breastfeeding, or within 3 months of giving birth. Structured written instruction in Arabic and English must be provided with every kit. 5.Single-use compliance must be enforced and monitored; reuse or sharing of any component is prohibited. Kits must be used within the labelled 24-month expiry, stored at 2–30 °C, and specimens must reach the laboratory within 7 days of collection. 6.All Kit expiries or malfunctions, deaths and serious injuries must be reported to DOH Pharmacovigilance at PVE@doh.gov.ae Clinical Use Overview •Population / Intended User — Adult women eligible for cervical cancer screening. •To be Performed By — Specimen collection is performed by the woman herself. •Clinical Setting — DOH-licensed healthcare facility or supervised community screening programme. •Condition of Use — Self-collection, transport and preservation of a cervico-vaginal specimen for laboratory human papillomavirus analysis; not for diagnosis and not for use as a stand-alone test. Contraindications and exclusions The Human Papillomavirus Self Sample Collection Kit is contraindicated for use in: •Women suffering from any gynecological disease •Women in the middle of a menstrual cycle, or with any vaginal bleeding or injury •Women who are pregnant, suspect they may be pregnant, or are breastfeeding •Women within 3 months of giving birth •Use of any other vaginal product within 2 days before collection, other than vaginal contraceptives, condoms and water-based lubricants •Any kit with damaged packaging or contents, a broken tube, or leaking preservation solution •Any kit used after the labelled expiry date, or any sampler with loose or shedding bristles •Reuse or sharing of any component — the kit is single use only General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26050 | LithiX™ Hertz Contact Intravascular Lithotripsy Catheter (LithiX HC-IVL) | Al Naghi Medical | Elixir Medical Corporation | USA | Medical Device | Therapeutic | Approved for adoption for calcium fragmentation of moderate to severely calcified, stenotic de novo coronary artery lesions prior to stenting. Recommendations: 1.The LithiX™ HC-IVL Catheter is recommended for use in licensed hospital cardiac catheterisation laboratories for calcium fragmentation in moderate-to-severely calcified, stenotic de novo coronary lesions prior to stent implantation, where clinically indicated. 2.Use must be restricted to interventional cardiologists trained and credentialled in angiography and percutaneous transluminal coronary angioplasty, as required by the Instructions for Use. Manufacturer-delivered training, proctored initial cases and documented competency sign-off are mandatory before first independent clinical use. 3.Facilities must have on-site or rapidly accessible emergency coronary artery bypass surgery and haemodynamic support capability. 4.Case selection must be documented by the treating interventional cardiologist, confirming a de novo lesion with moderate-to-severe calcification anticipated to resist full balloon dilation, and confirming the lesion is balloon-crossable. 5.Single-use compliance must be enforced and monitored; reuse, reprocessing or resterilisation is prohibited and carries a stated risk of patient injury, illness or death. Devices must be used within the labelled expiry and stored at ≤ 25 °C. 6.All device malfunctions, deaths and serious injuries must be reported to the DOH Pharmacovigilance at PVE@doh.gov.ae with balloon rupture, coronary perforation, dissection, slow or no-reflow and device entrapment tracked as device-specific events of interest. Clinical Use Overview •Population / Intended User — Adult patients undergoing percutaneous coronary intervention for moderate-to-severely calcified, stenotic de novo coronary artery lesions requiring lesion preparation prior to stent implantation, where the stenosis is anticipated to resist full balloon dilation and uniform stent expansion. Reference vessel diameter 2.25–3.5 mm and lesion length ≤ 34 mm per the studied evidence base. •To be Performed By — Interventional cardiologist trained and credentialled in angiography and percutaneous transluminal coronary angioplasty, supported by trained cardiac catheterisation laboratory nursing and radiographic staff. •Clinical Setting — Licensed hospital cardiac catheterisation laboratory with high-quality fluoroscopy, intravascular imaging (OCT/IVUS) capability, haemodynamic support availability, and on-site or rapidly accessible emergency coronary artery bypass surgery. No capital equipment, generator or console required. •Condition of Use: treating for calcium fragmentation of moderate to severely calcified, stenotic de novo coronary artery lesions prior to stenting. Contraindications and exclusions Limitations Safety and efficacy have not been evaluated in patients with: •Unresolved thrombus at the lesion site •Coronary artery spasm of the target vessel in the absence of a significant stenosis Contraindications The LithiX Hertz Contact Intravascular Lithotripsy Catheter is contraindicated for use in: •Patients judged to have a lesion that prevents complete inflation of an angioplasty balloon — uncrossable and undilatable disease •Patients with coronary artery spasm in the absence of significant stenosis •Post dilatation of expanded stent •Carotid or cerebrovascular arteries Additional exclusions •In-stent restenosis •Delivery through the side cell of a previously placed stent •Any target vessel where the required inflated balloon diameter would exceed the reference vessel diameter Clinical judgement required •Patients who are not acceptable candidates for coronary artery bypass graft surgery — treatment requires careful consideration including possible procedural haemodynamic support, as this population carries special risk •Patients in whom anticoagulation cannot be appropriately maintained during and after the procedure •Lesions requiring balloon sizes outside the 2.25–3.5 mm reference vessel diameter band studied in PINNACLE-I, i.e. the 1.5 mm, 2.0 mm and 3.25 mm sizes •Lesions exceeding the 14 mm balloon length, requiring sequential distal-then-proximal treatment General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26049 | PANA RealTyper HPV 32 Genotyping Kit | Babirus Medical Equipment Trading LLC | HLB Panagene Co., Ltd. | Republic of Korea | Diagnostic Test | Diagnostics | Approved The PANA RealTyperTM HPV 32 Genotyping kit is an in vitro diagnostic reagent for detecting human papilloma virus (HPV) using peptide nucleic acid (PNA) probes. This kit is a DNA amplification test for the qualitative detection of a total of 32 types of HPV by using a real-time PCR (polymerase chain reaction) system. This kit enables target classification by type (20 high-risk and 12 low-risk types) based on melting temperature (Tm) analysis. Detectable genotypes by PANA RealTyperTM HPV 32 Genotyping kit : [ High-risk genotypes (20 genotypes) ] - 16, 18, 26, 31, 33, 35, 39, 45, 51, 52, 53, 56, 58, 59, 66, 68, 69, 70, 73, and 82 [ Low-risk genotypes (12 genotypes) ] - 6, 11, 32, 34, 40, 42, 43, 44, 54, 61, 72 and 81 Recommendation: * Adopt the assay for appropriate clinical use: The PANA RealTyper™ HPV 32 Genotyping Kit should be implemented in clinical molecular laboratories for women undergoing cervical cancer screening or diagnostic evaluation for suspected HPV infection. Test results should be interpreted together with cervical cytology, clinical findings, and other diagnostic procedures when indicated. * Ensure laboratory verification and personnel training: Before routine implementation, laboratories should verify assay performance under local operating conditions, including analytical accuracy, precision, calibration, and quality control. Laboratory personnel should receive appropriate training in molecular diagnostic techniques and interpretation of HPV genotyping results. * Maintain quality assurance and regulatory compliance: Laboratories should follow established quality management procedures, including the use of recommended instruments and consumables, routine quality control, reagent storage according to the manufacturer’s instructions, and participation in external quality assessment programs where available. * Monitor emerging clinical evidence and screening recommendations: As cervical cancer screening guidelines continue to evolve, healthcare providers and laboratories should remain updated regarding HPV screening recommendations, genotype-specific clinical significance, and new evidence supporting HPV molecular diagnostic technologies. Clinical Overview: Population Women undergoing routine cervical cancer screening or diagnostic evaluation for suspected HPV infection. To be performed by Qualified clinical laboratory professionals trained in molecular diagnostic testing. Clinical setting Clinical molecular diagnostic laboratories, hospitals, and accredited healthcare facilities. Condition of use In vitro molecular diagnostic assay performed on cervical clinical specimens for HPV DNA detection and genotyping. Results should be interpreted in conjunction with cervical cytology, clinical findings, and other diagnostic procedures, and should not be considered a standalone diagnosis of cervical cancer. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26047 | Transonic FlowXL | Bayan Medical Company | TRANSONIC | USA | Medical Device | Therapeutic | Approved for use with adult and pediatric patients for the measurement of blood and liquid volume flow as follows: • on major and peripheral arteries, veins, and ducts, • on non-aerated synthetic vessel grafts. • during open surgical procedures where surgery has already been medically indicated, • at intraoperative sites which admit and retain ultrasonic couplant, such as sterile saline/ultrasonic gel • couplant, • where their application requires minimal vessel manipulation or constriction which could cause vessel • spasm that would interfere with volume flow measurements, • where their use does not unnecessarily lengthen surgical procedures, Tubing flow measurements with the FlowXL and Flowsensors are indicated for use on flexible tubing during extracorporeal applications, (not on arteries or veins), and only for non-aerated media which is transparent to ultrasound. Recommendations · Adopt FlowXL as an adjunctive tool for intraoperative graft-flow assessment during CABG surgery. · Use the device together with clinical judgement, visual assessment and other standard graft-evaluation methods. · Provide formal training for surgeons, nurses and perfusionists on probe selection, placement, signal verification and interpretation of flow measurements. · Establish clear procedures for cleaning, sterilizing and inspecting reusable Flow probes before each use. · Minimize electromagnetic interference and confirm abnormal readings using alternative assessment methods when necessary. · Monitor clinical use, graft revisions, device-related incidents and maintenance requirements after implementation. Clinical Use Overview: Population / Intended User: Perivascular Flow probes are used in adult or pediatric patients during surgical interventions where surgery is medically indicated and where an open surgical field allows Flow probe access to the vessels being repaired, augmented, or restricted. Tubing Flowsensors are used in adult or pediatric patients that are undergoing extracorporeal applications where flow measurements could be of use in evaluating fluid delivery or removal, such as heart/lung bypass, extracorporeal membrane oxygenation, or other tubing flows such as infusions. To be Performed By: The use of the FlowXL, Flow probes, and Flowsensors is restricted to medical professionals/clinicians. Disclaimer: “Trained clinicians” in the context of using the DUE (device under evaluation) refers to healthcare professionals who have undergone training in the fundamental surgical methods and procedures relevant to the application of the DUE. This implies that while they may not be specifically trained in using the DUE, they possess the necessary expertise in the underlying surgical techniques and procedures during which the DUE are employed. Clinical Setting: The FlowXL, Flow probes, and Flowsensors are intended for use in a clinical environment (e.g., an operating room). Condition of Use: installed and maintained according to IFU requirements. Contraindications: There are no Contraindications for use with the FlowXL, Flow probes or Flowsensors, however there are certain non-use cases. For these and Warnings and Precautions, please read the following sections carefully. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26046 | QT Scanner 2000 Breast Acoustic CT | Al Naghi Medical | QT Imaging Holdings, Inc. / QT Imaging, Inc. | United States of America | Medical Device | Diagnostics | Conditionally approved for adoption as an adjunctive breast imaging modality for use as an ultrasonic imaging system to provide reflection-mode and transmission-mode images of a patient's breast. The QT Scanner software also calculates the breast fibroglandular tissue volume (FGV) value and the ratio of FGV to total breast volume (TBV) value as determined from reflection-mode and transmission-mode ultrasound images of a patient's breast. The device is not intended to be used as a replacement for screening mammography. The QT Scanner is indicated for use by trained healthcare professionals in environments where healthcare is provided to enable breast imaging in adult patients. Recommendations: 1.QT Scanner 2000 Model B Breast Acoustic CT is recommended for hospital / radiology use as an adjunctive breast imaging tool in adult patients, according to the IFU indication. 2.It must not be used as a replacement for screening mammography. 3.To be operated only by trained and competent healthcare professionals after manufacturer or authorized distributor training. 4.Radiologist interpretation must be performed with correlation to clinical history, mammography/DBT, ultrasound, MRI, pathology, and BI-RADS pathway where applicable. 5.The facility must complete site readiness before installation, including room size, electrical supply, DI water, drainage, network, PACS/DICOM archiving, and cybersecurity/data privacy requirements. 6.DICOM images must be archived according to local retention requirements; RAW data backup and any AWS transfer must follow UAE data protection and health ICT requirements. 7.Infection prevention measures must be followed, including correct chlorine concentration, cleaning between patients, and single-use adhesive breast retention device tape. 8.Do not scan patients with contraindications listed below. 9.Quality control phantom scans, preventive maintenance, and acceptance testing must be performed according to QT Imaging requirements before patient use. 10.Clinical outcomes, recall rates, image quality issues, incomplete scans, adverse events, malfunctions, and user complaints should be monitored and reported as part of post-market surveillance. Clinical Use Overview: •Population / Intended User: Adult patients requiring breast imaging; intended users are trained healthcare professionals. •To be Performed By: Trained radiology technologists / trained operators; interpreted by qualified radiologists. •Clinical Setting: Licensed healthcare facility, radiology department, breast imaging center, or imaging suite. •Condition of Use: Prescription/professional use only; controlled clinical environment; installed and maintained according to IFU and pre-installation requirements. •Contraindications: · Not for imaging tissues other than breast tissue. · Do not use in patients advised not to immerse the breast in water, including post-biopsy patients. · Do not use in patients with open wounds or sores on the breast. · Do not use in pregnant or lactating patients. · Do not use in patients allergic to chlorine or alcohol wipes. · Do not use in patients allergic to medical-grade adhesive tape. · Do not scan patients with magnetic implants contraindicated for MRI, such as pacemakers, because the breast retention device and retention rod use magnets. · Do not use if patient weight exceeds 181 kg / 400 lb. · Do not use if breast size does not fit comfortably through the provided table insert ring, approximately outside A to DDD cup range. AAMEN Decision : Conditionally Approved till 12 February 2027 -Data shall be anonymized within the UAE prior to any transfer outside the UAE. -Access to health data and PHI/PII shall be restricted to authorized personnel only. -L1 and L2 support services shall be provided from within the UAE. -The applicant shall ensure ongoing alignment with applicable UAE laws and regulations. -Right to audit by regulatory body (DOH). General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26045 | Visualase™ MRI-guided laser ablation System | Bayan Medical Company LLC | Medtronic | USA | Medical Device | Therapeutic | Approved for The Visualase™ MRI-Guided Laser Ablation System is indicated for use as a minimally invasive procedure to ablate, necrotize or coagulate soft tissue through interstitial irradiation or thermal therapy under magnetic resonance imaging (MRI) guidance for wavelengths 800nm through 1064nm in cranial neurosurgery. Recommendation: •Use only in specialist neurosurgical centers. •MDT case selection required. •Restricted to patients with clearly defined intracranial target lesions. •consent explaining evidence limitations and neurological risks. •audit/registry collection for outcomes and adverse events. AAMEN Conditions: Support staff (L1, L2 and L3) must not have access to PHI or PII. Clinical Overview: Population: Adult patients with brain tumors, radiation necrosis, or drug-resistant focal epilepsy who are suitable candidates for MRI-guided laser ablation. To be performed by: Performed by a trained neurosurgeon experienced in MRI-guided laser interstitial thermal therapy Clinical setting: Healthcare facilities with neurosurgery and MRI facilities. Condition of use: Used under MRI guidance for carefully selected patients after multidisciplinary team evaluation, particularly when open surgery is not feasible or carries a high risk. Contraindications: As per user manual: We strongly recommend physicians weigh advantages and disadvantages of using a diode laser. Other modalities or wavelengths may be more appropriate due to any of the following: •Depth of penetration •Volume of necrosis •Propensity for scarring This product should not be used if thermal therapy or interstitial laser therapy are contraindicated. Use the laser only for the specialties listed in the Indications for Use section. Do not use the device endoscopically in any procedure where endoscopes are contraindicated. Patients who are unable to be treated by surgical means or who are intolerant to anesthesia. General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-09-09 | HTA26040 | Magstim TMS Therapy -Horizon® 3.0 with StimGuide | BAYAN MEDICAL COMPANY | The Magstim Company Ltd. | United Kingdom (UK) | Medical Device | Therapeutic | Approved for adoption for the treatment of Major Depressive Disorder in adult patients who have failed to achieve satisfactory improvement from prior antidepressant medication in the current episode, and as an adjunct for the treatment of adult patients suffering from Obsessive-Compulsive Disorder. Recommendations: 1.The Magstim® Horizon® 3.0 with StimGuide® Pro TMS Therapy System is recommended for use in licensed hospitals or outpatient psychiatric/neuromodulation clinics for repetitive transcranial magnetic stimulation of adult patients with treatment-resistant Major Depressive Disorder, and as an adjunct in adult Obsessive-Compulsive Disorder, where clinically indicated. 2.Use must be restricted to licensed physicians (psychiatrists) trained in TMS, with delivery by trained technicians or nurses under documented physician supervision; manufacturer-delivered training and competency sign-off is mandatory before first clinical use. 3.The system must be used only within its FDA-cleared indications and parameter set: 1.standard rTMS — 10 Hz, 120% motor threshold, 4-second trains with 11–26 second inter-train intervals, 75 trains, 3,000 pulses per session, 18.8–37.5 minutes, 5 sessions weekly for 6 weeks, left dorsolateral prefrontal cortex; 2.iTBS — 120% motor threshold, 50 Hz triplets at 5 Hz burst rate, 600 pulses per session, 3.09 minutes, 5 sessions weekly for 6 weeks, left DLPFC; 3.OCD adjunct — 20 Hz, 120% motor threshold, 2-second trains with 20-second inter-train intervals, 50 trains, 2,000 pulses per session, 18.3 minutes, 5 sessions weekly for 5 weeks followed by 4 sessions in the sixth week, dorsomedial prefrontal cortex. 4.Patient selection must be documented by a qualified psychiatrist, confirming failure of at least one adequate antidepressant trial in the current episode (MDD), or confirming established SSRI and exposure-and-response-prevention therapy is in place and maintained at current doses (OCD). 5.A pre-treatment screening questionnaire must be completed for every patient, with all Operating Manual contraindications applied as hard stops in the electronic order set — including the 30 cm implant and ferromagnetic metal exclusion zone and the age limits of 22–70 years (MDD) and 22–68 years (OCD). 6.Motor threshold must be determined and documented by a competent operator at the start of each treatment course, and re-determined following any change in medication that could alter seizure threshold. 7.Hearing protection rated at 30 dB noise reduction must be worn by the patient in every session. 8.The facility must maintain a biomedical engineering electrical safety, EMC and preventive maintenance programme, including annual verification of the coil temperature interlock and calibration of the StimGuide Pro optical tracking system; no modification of the equipment is permitted and servicing is restricted to qualified personnel. 9.Seizures, hypomania or mania, syncope, thermal injury, and all serious adverse events or device malfunctions must be reported through post-market surveillance and to the DOH Pharmacovigilance PVE@doh.gov.ae 10.Treatment administrators must be trained to recognize and respond to seizure, and must stop treatment immediately on excessive facial, hand, leg or neck motor activation outlasting the pulse train. 11.A facility registry must capture indication, protocol, sessions delivered, motor threshold, targeting method, outcome scores and adverse events for every treated patient, with response and remission rates reported to DOH at 12 and 24 months and benchmarked against the class expectation of 58–83% response and 28–62% remission. Clinical Use Overview •Population / Intended User — Adults aged 22 to 70 years with Major Depressive Disorder who have failed to achieve satisfactory improvement from prior antidepressant medication in the current episode; and adults aged 22 to 68 years with Obsessive-Compulsive Disorder receiving adjunctive treatment. •To be Performed By — Licensed physician (psychiatrist) trained in TMS; delivery by a trained TMS technician or nurse under documented physician supervision. •Clinical Setting — Licensed hospital (inpatient or outpatient), outpatient psychiatric clinic, or neuromodulation clinic with on-site seizure management capability. No anaesthesia required. •Condition of Use — Prescription use only; one of the three cleared stimulation protocols at 120% of individually determined motor threshold; mandatory pre-treatment screening questionnaire; 30 dB hearing protection every session for patient and for operators within 2 metres; earth straps fitted as supplied; minimum 30 cm separation from portable RF communications equipment; equipment not to be stacked with or placed adjacent to other electromedical equipment. Contraindications: •Patients with implants activated or controlled by physiological signals in/near the head or within 30 cm of the coil — including inactive devices •Patients with device leads •Patients with cochlear implants •Patients with deep brain stimulators •Patients with vagus nerve stimulators •Patients with wearable cardioverter defibrillators — even if the device has been removed •Patients with conductive, ferromagnetic or magnetic-sensitive metal in the head or within 30 cm of the coil •Patients with stimulator devices, electrodes, aneurysm clips or coils, or stents •Patients with bullet fragments or metallic debris in the eye •Exemption: standard amalgam dental fillings are not contraindicated Operator and Bystander Exclusions •Pregnant women — including staff •Anyone with implanted medical devices •Anyone with implanted metal •Anyone wearing removable magnetically sensitive items — earrings, hearing aids, eyeglasses, cell phones, MP3/MP4 players, wearable monitors, bone growth stimulators Clinical Judgement Required •Patients with physiologically-controlled implants beyond 30 cm — pacemakers, ICDs, WCDs •Patients with non-physiologically-activated implants or metal beyond 30 cm •Patients with epilepsy, seizure history or family history of seizures •Patients with brain lesions affecting seizure threshold •Patients with stroke, head injury, cerebrovascular accident, cerebral aneurysm, dementia or movement disorders •Patients on tricyclic antidepressants, neuroleptics or other seizure-threshold-lowering drugs •Patients with sleep deprivation •Patients with heavy alcohol consumption or using epileptogenic drugs •Patients with severe heart disease •Patients with increased intracranial pressure •Patients with uncontrolled migraines •Patients with conditions altering electrolyte balance •Patients with no quantifiable motor threshold •Patients with diabetes mellitus (thermal injury risk) •Patients under the influence of alcohol or using sleep medication (thermal injury risk) •Patients with broken skin in the contact area •Pregnant patients Safety Not Established: •Patients younger than 22 years •Patients older than 70 years (MDD) or 68 years (OCD) •Patients with no prior antidepressant medication failure (MDD) •Patients with a suicide plan or recent suicide attempt •Patients on concurrent antidepressants unable to tolerate discontinuation •Patients with seasonal affective disorder •Patients with a history of substance abuse, OCD or PTSD (MDD); substance abuse in the previous six months (OCD) •Patients with depression secondary to a medical condition or substance-induced •Patients with psychotic disorder, schizoaffective disorder, bipolar disease or MDD with psychotic features •Patients with severe personality disorder or severe borderline personality disorder (OCD) •Patients who failed to benefit from ECT or VNS, or currently receiving these therapies •Patients with a history of increased intracranial pressure or head trauma •Patients who are pregnant or nursing •Patients with implanted medication pumps, intracardiac lines or significant cardiac disease (OCD) •Patients on high-dose psychotropics lowering seizure threshold, e.g. clomipramine (OCD) •Patients with psychoses or psychiatric emergencies requiring rapid response — catatonia, marked physical deterioration, immediate suicide risk General Requirements: •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA25098 | AiTiALVSD | Al Naghi Medical | Medical AI Co., Ltd. | Korea | System/ Software/ AI/ App | Prevention | Approved: AiTiALVSD is software that utilizes AI algorithms to support diagnosis. AiTiALVSD is designed to assess LVSD (≤40%) in adults at risk of heart failure using 12-lead ECG data. • Adoption: Consider incorporating AiTiALVSD in adult primary care and cardiology ECG workflows to enhance early HF screening. Educate clinicians on interpreting the AI risk score and reinforce that it complements (not replaces) clinical evaluation. Begin with targeted use (e.g. patients with risk factors) as experience is gained. • Verification & Implementation: Validate the software on local ECG/Echo data if possible, and train staff. Establish a clinical pathway: e.g., automatic follow-up echo for high-risk scores. Ensure IT integration and uptime. • Quality Control & Ongoing Assessment: Monitor key performance metrics (e.g. proportion of AI-flagged patients who receive echo and true positive rate). Periodically audit AI outputs versus echo results to detect drift or bias. Update and recalibrate the algorithm as new data become available. Ensure maintenance of data security and regulatory compliance. AAMEN Approved: -Data residency sovereignty. -No PHI/PII shall be transferred or accessible outside the UAE borders. -Support service sovereignty. -Support (L1, L2 and L3) shall have no visibility of PHI/PII. -L1 and L2 support shall be within UAE. -Right to Audit by regulatory body. Clinical Overview: Population: Adult patients (≥18 years) undergo ECG without known HF. Performed by: medically trained health care professionals. Clinical Setting: Healthcare facilities Condition of Use: Run on each eligible 12 lead ECG as a screening step; if AiTiALVSD flags high risk, order a transthoracic echocardiogram to confirm LVSD. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26007 | Cercare Medical Neurosuite | Emerging Global Technologies | Cercare Medical A/S | Denmark | System/ Software/ AI/ App | Others | Approved for Adoption: 1. Approved for adoption for use as image processing software package to be used by trained professionals, including physicians and medical technicians.The software runs on standard off-the-shelf computers or a virtual platform, such as VMware, and can be used to perform image viewing, processing and analysis of medical images. Data and images are acquired through DICOM compliant imaging devices. CMN provides viewing capabilities for all MR and CT DICOM images of single-frame or enhanced types. 2. To be used as an adjunctive medical image processing and decision-support software, not as an autonomous diagnostic tool. Recommendations: 1. Cercare Medical Neurosuite is recommended for hospital use as adjunctive medical image processing and decision-support software. 2. To be used only by trained healthcare professionals with radiology/stroke workflow knowledge. 3. Software outputs must be reviewed and confirmed by clinicians/radiologists before clinical decisions. 4. The software must not replace standard CT/MRI interpretation, PACS review, radiology reporting, or clinical judgment. 5. LVO, ICH, ASPECTS, perfusion, and mismatch outputs should be used only within their approved intended use. 6. Integration with PACS/RIS/DICOM workflow should be validated locally before clinical go-live. 7. Data privacy, cybersecurity, ADHICS, and UAE health data hosting/transfer requirements must be ensured. 8. All AAMEN requirements must be fulfilled based on the approval; • Support (L1, L2 and L3) shall have no access or visibility of PHI/PII. • L1 and L2 support shall be within UAE. • Patient Data shall not be used to train the AI/ML model. • System logs shall be made available for the health entity. 9. False positive, false negative, delayed alert, and workflow interruption events should be monitored through post-market surveillance. 10. Periodic user training, competency validation, and software performance review should be implemented. Clinical Use Overview: • Population / Intended User: General imaging population for CMN/CF; adult suspected acute stroke patients for VE, ICH, and LVO modules. • To be Performed By: Radiologists, neurologists, stroke physicians, trained physicians, and medical technicians. • Clinical Setting: Hospitals, radiology departments, stroke units, emergency departments, and healthcare research centers. • Condition of Use: Adjunctive decision support only; final diagnosis and treatment decisions remain with the responsible clinician. • Contraindications: No direct software contraindications. Imaging-device and contrast-agent contraindications, such as MRI-unsafe implants or contrast restrictions, must be managed during image acquisition. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26025 | Fully Automated Molecular POCT Analyzer - HysenDx 200 / DX200 | Omics One Sciences Medical Supplies Trading LLC | Hysen Biotech.Inc | Republic of Korea / Korea (South) | Diagnostic Test | Diagnostics | Approved for :adoption and used as quantitative detection of analytes in the related pathogen nucleic acids (DNA/RNA) from human samples using the polymerase chain reaction process. The samples can be nasopharyngeal and oropharyngeal samples, whole blood, blood plasma, blood serum, saliva samples and so on. The instrument is for in vitro diagnostic only. The Fully Automated Molecular POCT Analyzer is intended for used in medical and biological laboratories by professional user trained in molecular biological techniques and the operation of Fully Automated Molecular POCT Analyzer. Recommendation: 1. Hysen Fully Automated Molecular POCT Analyzer (HysenDx 200) is recommended for restricted hospital laboratory use for rapid molecular detection of pathogen nucleic acids, in line with the IFU and validated compatible assays only. 2. To be operated only by trained and competent healthcare laboratory professionals experienced in molecular testing and biosafety procedures. 3. Use only manufacturer-provided or manufacturer-recommended consumables, chips, and accessories; reuse of chips/consumables is prohibited. 4. The device should be installed and used only in appropriate indoor laboratory clinical environments with proper grounding, ventilation, spacing, and temperature/humidity controls. 5. Operators must observe all biohazard and hot-surface precautions, including PPE, avoidance of touching heated modules, and adherence to PCR laboratory regulations. 6. The analyzer should be integrated into infectious disease - respiratory testing pathways where rapid molecular results can influence isolation, cohorting, treatment selection, and antimicrobial stewardship. 7. Initial installation, annual recalibration, maintenance, and troubleshooting should be supported by the manufacturer or authorized service personnel only. 8. Post-market monitoring should include documentation of invalid-failed runs, turnaround time, concordance with central laboratory testing, contamination incidents, maintenance issues, and user-reported adverse events/malfunctions. Clinical Use Overview: A rapid, near-patient or decentralized molecular PCR analyzer for infectious disease diagnosis, especially respiratory pathogen detection when a fast sample-to-result workflow is needed. Population / Intended User: Patients suspected of infectious disease; end users are trained laboratory/healthcare professionals. To be Performed By: Trained professional users in medical and biological laboratories; practically this includes laboratory personnel and other trained healthcare professionals according to institutional policy. Clinical Setting: Hospitals, medical laboratories, biological laboratories, and suitable decentralized healthcare settings where molecular POCT is operationally appropriate. Condition of Use: The Fully Automated Molecular POCT Analyzer is an automated instrument used for quantitative detection of analytes in the related pathogen nucleic acids (DNA/RNA) from human samples using the polymerase chain reaction process. The samples can be nasopharyngeal and oropharyngeal samples, whole blood, blood plasma, blood serum, saliva samples and so on. The instrument is for in vitro diagnostic only. The Fully Automated Molecular POCT Analyzer is intended for used in medical and biological laboratories by professional user trained in molecular biological techniques and the operation of Fully Automated Molecular POCT Analyzer. Contraindications: No explicit patient contraindications. The main restrictions are operational: not for use outside intended purpose, not with non-approved consumables, not in hazardous environments, and not with prohibited flammable/explosive/reactive materials in the centrifuge component. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26027 | Oculera Visual Field Analyzer (Oculera VFA) | Oculera Vision Tech LLC-FZ | Oculera Sağlık Teknolojileri A.Ş. | Turkey / Türkiye | Medical Device | Diagnostics | Approved for adoption as an automated visual field assessment under supervision of licensed healthcare professionals for assessment/monitoring of visual field defects associated with glaucoma and neuro-ophthalmic disorders, and in school as a screening/referral support tool for children aged 6 years and above. This adoption statement should be interpreted in line with the uploaded IFU, EDE classification conditions, and the available EU MDR Class IIa/QMS documentation. Recommendations: • Oculera Visual Field Analyzer 2.8.0 is recommended for controlled clinical use for automated visual field assessment under the supervision of licensed healthcare professionals, and for school-based screening/referral support only within the stated intended use. • Not diagnostic testing • To be operated only by trained and competent personnel, with clinical interpretation performed by qualified clinicians. • The technology should not be used as a replacement for comprehensive ophthalmic examination or as a stand-alone diagnostic decision tool in school settings. • Screening/testing should be discontinued immediately if the patient experiences dizziness, nausea, distress, anxiety, or other VR-related adverse symptoms. • Proper headset fitting, fixation control, environment preparation, and patient instruction must be ensured to optimize test reliability. • Only approved software versions and authorized system configurations should be used; unauthorized modification is not permitted. • The device should not be used in children under 6 years, in severe low vision (≤2/10), active ocular infection, significant facial/skin infection in the contact area, or in users unable to cooperate with instructions. • Facilities adopting the system should monitor real-world performance, referral yield, reliability indices, adverse events, and repeat-test consistency as part of routine post-market surveillance and local quality reporting. This is especially important because the current evidence base is still limited. AAMEN Approved: - No PHI/PII shall be transferred or accessible outside the UAE borders. - Support (L1, L2 and L3) shall have no visibility of PII. - L1 and L2 support must be available within the UAE. Clinical Use Overview: Oculera VFA is a non-invasive VR-based visual field assessment and screening platform intended for supervised use in ophthalmic and screening workflows. It is portable and designed to function without a dark room, which may improve usability in clinics and school-based programs. Population / Intended User: Patients ≥ 6 years who require visual field testing and can cooperate with instructions; in school deployment, school-aged children ≥6 years for screening/referral support. To be Performed By: Medical staff / trained and authorized operators under established protocols; interpretation by licensed healthcare professionals. Clinical Setting: Ophthalmology clinics, healthcare facilities, and school screening environments using the Smart Screening deployment model. Condition of Use: Under supervision of licensed healthcare professionals, with secure configuration, correct headset positioning, controlled environment, and appropriate patient instruction. Contraindications: Age under 6 years Severe low vision (≤2/10) Active ocular infection Significant facial or skin infection/condition in areas contacting the headset Inability to cooperate with instructions Clinician judgment may expand or restrict use case-by-case. One important note for the file record: the submission contains strong documentation for EDE classification and EU MDR Class IIa/QMS status, but I did not find uploaded official FDA or MHRA certificates, so I treated those as applicant-stated, not independently verified. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26031 | Advanced Hybrid Closed Loop MiniMed™ 780G Insulin Pump System | Modern Pharmaceutical Company L.L.C | Medtronic | USA | Medical Device | Therapeutic | Approved as The MiniMed 780G system is indicated for use by individuals age 2 years and older with insulin-requiring diabetes, whose total daily dose of insulin is 6 units per day or more. • Use only manufacturer-approved MiniMed™ 780G components, including the Guardian™ 4 sensor, Guardian™ 4 transmitter, compatible Medtronic infusion sets, reservoirs, and approved U-100 insulin. Mixing incompatible components is not recommended. • The system should be initiated only after structured education by a qualified diabetes healthcare professional, including pump operation, carbohydrate counting, infusion-set management, alarm response, sick-day management, ketone monitoring, and emergency backup insulin administration. • SmartGuard automated insulin delivery should be used only within the approved operating conditions and recommended daily insulin requirements (6–250 units/day), following the manufacturer’s instructions for calibration-free operation, maintenance, software updates, and replacement of consumables. • Blood glucose meter measurements should be used whenever sensor glucose readings do not match the patient’s symptoms or during situations known to interfere with sensor accuracy (e.g., certain medications), and clinical decisions should not rely solely on CGM values when discrepancies exist. • The pump, sensor, and transmitter should be removed before MRI, CT, X-ray, diathermy, or other procedures involving strong magnetic fields or radiation, in accordance with the manufacturer’s instructions. • Patients should maintain an emergency backup plan, including rapid-acting insulin, blood glucose and ketone testing supplies, spare infusion sets, reservoirs, batteries, and an alternative insulin delivery method in case of device failure. • Post-market monitoring should include documentation of severe hypoglycaemia, diabetic ketoacidosis, hyperglycaemia due to infusion-set failure, device malfunctions, sensor failures, communication errors, software or connectivity issues, user complaints, recalls, adverse events, and patient-reported outcomes such as time in range, HbA1c, and treatment satisfaction. Clinical Use Overview: Population / Intended User: children aged ≥2 years, adolescents and adults with insulin-requiring diabetes using ≥6 units/day To be Performed By: patient or trained parent/caregiver with initiation and oversight by a diabetes specialist team. Clinical Setting: home environment Condition of Use: only after structured training, compatible components, individualized settings and continuous monitoring. AAMEN Exemption until 31st December 2026: 1. Access to health data and PII shall be restricted to authorized personnel only. 2. Obtain DoH web and mobile application approval. 3. Right to audit by regulatory body (DOH) General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26037 | GENEMATRIX NeoPlex HPV29 Detection (50, 96 Test) - Model NS02A | Omics One Sciences Medical Supplies Trading LLC | GeneMatrix Inc. | Republic of Korea | Medical Product/ Kit | Diagnostics | Approved for adoption for professional laboratory use for HPV29 Detection, a multiplex real-time PCR assay, was evaluated for the simultaneous qualitative detection of 29 HPV genotypes from liquid-based cytology specimens. Recommendations: 1.NeoPlex™ HPV29 Detection is recommended for use in licensed hospital or reference laboratories for qualitative detection and confirmation of HPV 29 genotypes from liquid-based cytology specimens where HPV testing/genotyping is clinically indicated. 2.Use must be restricted to trained and competent laboratory professionals only. 3.The test must be used only for the approved specimen type: liquid-based cytology specimens. 4.Results must be interpreted by qualified healthcare professionals within an approved cervical cancer screening or diagnostic pathway. 5.Positive HPV16/18 or other high-risk genotype results must be managed according to applicable cervical screening and colposcopy guidelines. 6.The kit must be stored and handled according to manufacturer requirements, including frozen storage at the specified condition and avoiding more than four freeze-thaw cycles. 7.Laboratories must follow contamination-control procedures, including unidirectional workflow, PPE, separate pre- and post-PCR areas, and appropriate positive/negative/internal controls. 8.The kit should not be used as a standalone diagnosis of cervical cancer; it detects HPV DNA and must be correlated with cytology, clinical findings, and guideline-based follow-up. 9.Local validation/verification should be completed before routine implementation, including verification on the laboratory’s PCR platform, specimen workflow, and reporting format. 10.Post-market monitoring should include invalid rate, false-positive/false-negative investigations, discordant results, lot performance, adverse events/malfunctions, and user complaints. 11.Any serious incident, malfunction, or safety concern must be reported through the responsible PV contact and applicable DoH reporting channels. Clinical Use Overview Population / Intended User: Women suspected of HPV infection or undergoing cervical screening/HPV genotyping. To be Performed By: Trained laboratory technologists / molecular diagnostics laboratory professionals. Clinical Setting: Licensed hospital laboratory, molecular diagnostics laboratory, or reference laboratory. Condition of Use: Professional in vitro diagnostic use only; liquid-based cytology specimen; use with validated nucleic acid extraction and real-time PCR workflow. Contraindications: • No direct patient contraindications, as the device is an in vitro diagnostic (IVD) assay. • Intended for professional use only by trained laboratory personnel. • Not intended for use with specimen types other than liquid-based cytology specimens. • Do not use specimens with insufficient residual volume (<200 µL). • Do not use specimens exposed for more than 24 hours without refrigeration (2–8°C). • Do not use specimens that have been frozen and thawed more than two times. • Do not use contaminated or improperly stored/transported specimens. • Do not use expired, contaminated, or mixed reagents from different lots. • Do not use the kit outside the approved Instructions for Use (IFU). • The clinical validation study included specimens from women aged 20–80 years; specimens outside this age range were excluded from the validation study. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26038 | PATHFAST D-Dimer | Al Shomoukh Trading for Technical & Medical Supplies – ASTEMED L.L.C., | PHC Europe | Japan | Diagnostic Test | Diagnostics | Approved as PATHFAST D-Dimer is a product for in-vitro diagnostic use with the in vitro diagnostic (IVD) automated analyser PATHFAST for the quantitative measurement of D-dimer in whole blood or plasma. PATHFAST D-Dimer is intended to be used: - as an aid in the diagnosis of activation processes of the coagulation system including deep vein thrombosis (DVT) and pulmonary embolism (PE), - by laboratory technician, nurse or physician, - in hospital including emergency room, doctor’s office and clinical laboratory. PATHFAST D-Dimer is a device for near patient testing (NPT). • Use only manufacturer-approved PATHFAST reagents, cartridges, calibrators, and consumables; reuse of cartridges or consumables is prohibited. • The device should be installed and operated in appropriate clinical environments according to the manufacturer’s instructions, including proper storage, calibration, and temperature conditions. • Results should always be interpreted together with patient symptoms, clinical assessment, and not used as the sole basis for diagnosis. • Calibration and quality control should follow the manufacturer’s recommendations, including regular user calibration and routine QC testing with every new reagent lot. • Post-market monitoring should include documentation of invalid runs, turnaround time, discordant clinical results, maintenance issues, recalls, and user-reported malfunctions or adverse events. Clinical Use Overview: Population / Intended User: Patients with suspected DVT/PE To be Performed By: by laboratory technician, nurse or physician Clinical Setting: in hospital including emergency room, doctor’s office and clinical laboratory. PATHFAST D-Dimer is a devance for near patient testing (NPT). Condition of Use: Use as a adjunct rule-out test, not as a standalone diagnostic test General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26039 | PATHFAST NTproBNP | Al Shomoukh Trading for Technical & Medical Supplies – ASTEMED L.L.C | PHC Europe | Japan | Diagnostic Test | Diagnostics | Approved as PATHFAST NTproBNP is a product for in-vitro diagnostic use with the in vitro diagnostic (IVD) automated analyser PATHFAST for the quantitative measurement of N-terminal pro-B-type natriuretic peptide (NT-proBNP) in human heparinized or EDTA whole blood and plasma. PATHFAST NTproBNP is intended to be used: - as an aid in the diagnosis in patients suspected of congestive heart failure (CHF) and detection of left ventricular dysfunction, - as an aid in the severity assessment of CHF, - in the risk stratification in patients with acute coronary syndromes (ACS) and CHF, - for monitoring the therapy in patients with left ventricular dysfunction, - by laboratory technician, nurse or physician, - in hospital including emergency room, doctor’s office and clinical laboratory. PATHFAST NTproBNP is a device for near patient testing (NPT). • Use only manufacturer-approved PATHFAST reagents, cartridges, calibrators, and consumables; reuse of cartridges or consumables is prohibited. • The device should be installed and operated in appropriate clinical environments according to the manufacturer’s instructions, including proper storage, calibration, and temperature conditions. • Results should always be interpreted together with patient symptoms, clinical assessment, and not used as the sole basis for diagnosis. • Calibration and quality control should follow the manufacturer’s recommendations, including regular user calibration and routine QC testing with every new reagent lot. • Post-market monitoring should include documentation of invalid runs, turnaround time, discordant clinical results, maintenance issues, recalls, and user-reported malfunctions or adverse events. Clinical Use Overview: Population / Intended User: Patients with suspected or established heart failure, in conjunction with clinical assessment and other diagnostic investigations. To be Performed By: by laboratory technician, nurse or physician Clinical Setting: in hospital including emergency room, doctor’s office and clinical laboratory. PATHFAST NTproBNP is a device for near patient testing (NPT). Condition of Use: Use as a adjunct rule-out test, not as a standalone diagnostic test General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26041 | GT Metabolic MagDI System | Al Naghi Medical | GT Metabolic Solutions, Inc. | USA | Medical Device | Therapeutic | Approved for The GT Metabolic MagDITM System is intended for use in the creation of side-to-side duodeno-ileal anastomoses in minimally invasive and laparoscopic surgery. Once wound strength is sufficient to maintain the anastomosis, the device is passed from the body. The effects of this device on weight loss were not studied. The GT Metabolic MagDITM System is intended for use in adult patients > 21 years. •Confirm MRI risk and implanted device status before use. •Maintain local registry/audit: technical success, device passage time, complications, re-intervention, anastomotic leak/bleeding/infection/obstruction. Population: adult patients > 21 years. To be performed by: Healthcare professional. Trained , gastroenterologists /surgeons Clinical setting: Healthcare facilities./ Used in the operating room during minimally invasive or laparoscopic gastrointestinal surgery. Condition of use: Intended for adult patients (>21 years) Contraindications: •Do not use the devices if the patient is not indicated for a side-to-side anastomosis of two (2) segments of the small bowel. •Do not use the devices if an immediate lumen patency is required while the anastomosis is formed. •Do not use the devices if the patient has unhealed ulcers, bleeding lesions, or a tumor or any other lesion at the target Magnet sites in the small bowel. •Do not use the devices if the patient has an expected need for magnetic resonance imaging (MRI) within 30 days after placement of the Magnets or until passage of the devices out of the body. The Magnets are MRI unsafe. •Do not use the devices if the patient has known allergies to the Magnet materials (Titanium alloy (ASTM F136), Parylene C, stainless steel, neodymium-iron-boron (ASTM A1101)) or the flange materials (polyglycolic-co-lactic acid (PGLA) or similar polymers, or barium sulfate). •Do not use the devices if the patient has an implanted pacemaker and/or defibrillator. •Do not use the devices if the patient has any other implanted electrical devices (e.g., neurological) or non-electrical implants or metal that may attract the Magnet devices. •Do not use the devices if the patient is pregnant or plans to become pregnant. •Do not use the devices if the patient has any conditions for which endoscopy or laparoscopic surgery would be contraindicated, and any significant congenital or acquired anomalies of the GI tract at or distal to the placement of the Magnets. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26042 | Cellvizio® I.V.E. and Confocal Miniprobes™ - FP-10001 | ALPHAMED GENERAL TRADING LL.C - SOLE PROPRIETORSHIP L.L.C. | Mauna Kea Technologies | France | Medical Device | Diagnostics | Approved for The Cellvizio® I.V.E. is a confocal laser system with fibre optic probes that are intended to allow imaging of the internal microstructure of tissues including, but not limited to, the identification of cells and vessels and their organisation or architecture. The Cellvizio® I.V.E. is intended to be used with a variety of Confocal Miniprobes™ which are provided with their own Instructions for Use • Use as an adjunct to standard diagnostic pathways • Follow the intended use and indication of the selected Confocal Miniprobe™. • Ensure operator training and competency before clinical use. • Conduct local monitoring of clinical outcomes and safety. Population: The intended patient population is people with a pathology or suspected pathology described in the specific I.V.E. Confocal Miniprobes™ User Guides. To be performed by: used by trained hospital personnel, qualified in the procedures for which it is intended, such as an operating physician, operating nurse, circulating nurse, or scrub nurse. Clinical setting: used under the supervision of an experienced operating physician in an endoscopy or an operating room during primary endoscopic, manual or robotic-assist laparoscopic, or interventional radiology procedures. Condition of use: Used during endoscopic, laparoscopic, robotic-assisted, or interventional procedures with compatible Confocal Miniprobes™. It is intended as an adjunct to standard diagnostic methods and not as a replacement for histopathology. Contraindication: There are no known contraindications associated with the use of the Cellvizio® I.V.E., provided the device is used according to the instructions provided in the User Guide. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26043 | VIASURE C. albicans, G. vaginalis & T. vaginalis Real Time PCR Detection Kit - Model: VS-CGT112H | OMICS ONE SCIENCES MEDICAL SUPPLIES TRADING L.L.C | CERTEST BIOTEC S.L. | Spain | Medical Product/ Kit | Diagnostics | Approved for adoption the specific identification and differentiation of Candida albicans, Gardnerella vaginalis and/or Trichomonas vaginalis in clinical specimens from patients with signs and symptoms of vaginal infection Recommendations: 1. VIASURE C. albicans, G. vaginalis & T. vaginalis Real Time PCR Detection Kit is recommended for use in licensed hospital or reference laboratories for detection and differentiation of C. albicans, G. vaginalis, and T. vaginalis DNA where clinically indicated. 2. Use must be restricted to trained and competent laboratory professionals experienced in real-time PCR and IVD procedures. 3. The assay should be used only within its approved intended purpose and validated specimen types: DNA extracted from vaginal and urethral swabs. 4. Results must be interpreted by qualified healthcare professionals together with symptoms, medical history, epidemiological risk factors, and other diagnostic findings. 5. Positive C. albicans or G. vaginalis results must not be used alone to confirm vulvovaginal candidiasis or bacterial vaginosis. 6. Negative results must not be used alone to exclude infection or guide treatment decisions. 7. Positive and negative controls, and internal/extraction controls, must be included and valid in each run. 8. Laboratories must follow contamination-control workflow, including separate extraction and amplification areas. 9. Specimen collection, transport, storage, and DNA extraction must follow the IFU and local laboratory procedures. 10. Invalid, ambiguous, or clinically discordant results should be repeated or confirmed using an alternative method. 11. Adverse incidents, serious malfunctions, invalid run trends, false-result concerns, and complaints should be reported through post-market surveillance and PV channels. AAMEN Approved : - PHI and PII are not transferred outside UAE. - Support from outside UAE don’t have access to PII and PHI - Right to audit by regulatory body (DoH). Clinical Use Overview: Population / Intended User: Patients with signs and symptoms of vaginal infection or suspected genitourinary infection/dysbiosis. To be Performed By: Qualified and trained clinical laboratory personnel. Clinical Setting: Licensed hospital laboratory, molecular diagnostics laboratory, reference laboratory, or clinical microbiology laboratory. Condition of Use: Professional IVD use only; validated PCR platform and extraction workflow; interpretation with clinical context. Contraindications: · No direct patient contraindications because this is an IVD test. · Not for use outside the validated intended purpose. · Not validated as a standalone diagnostic test for bacterial vaginosis. · Not for unvalidated specimen types unless locally validated. · Not for use by untrained personnel. · Do not use expired, damaged, improperly stored, or contaminated reagents. · Do not use if specimen collection, transport, storage, or extraction is inadequate. · Do not use results as the sole basis for treatment or patient management. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26044 | Edison System - Edison System version 3.x / Software version 3.1 | Al Naghi Medical | HistoSonics, Inc. | United States of America | Medical Device | Therapeutic | Approved for adoption as indicated for the non-invasive destruction of liver tumors, including unresectable liver tumors, using a non-thermal, mechanical process of focused ultrasound. Recommendations: 1. Edison System is recommended for hospital use for non-invasive destruction of liver tumors, including unresectable liver tumors, in accordance with FDA-cleared / IFU indication. 2. Use must be restricted to trained and competent physicians who completed HistoSonics formal training. 3. Case selection must be reviewed by a multidisciplinary tumor board, including oncology, hepatobiliary surgery, interventional radiology, hepatology, and anesthesia. 4. Patients must have adequate functional liver reserve before treatment. 5. Patients with Child-Pugh class C or decompensated cirrhosis should not be considered candidates. 6. Treatment should be performed in an appropriately equipped hospital procedure room / interventional radiology or operating room setting with anesthesia support. 7. Pre-procedure planning must confirm tumor visibility, acoustic pathway feasibility, treatment depth, and absence of critical structures within the treatment path. 8. A maximum of three hepatic tumors may be targeted in a single procedure. 9. The size of the targeted tumors must not exceed the device’s approved Instructions for Use (IFU) limits.10. Real-time ultrasound visualization of the bubble cloud must be maintained throughout treatment. 10. Treatment must be stopped if the patient, bed, treatment cart, or bubble cloud position moves or if PTV alignment is lost. 12. Facilities must comply with electrical safety, EMC, cybersecurity, PACS/DICOM, data privacy, cleaning, and maintenance requirements. 13. Clinical outcomes must be monitored, including technical success, complications, liver function, AKI, local tumor control, retreatment, local tumor progression, survival, and readmissions. 14. Adoption should include post-market surveillance and real-world registry reporting to confirm safety and effectiveness in Abu Dhabi practice. 15. As per AAMEN approval to ensure the following requirements: • PHI and PII are not transferred outside UAE. • Support from outside UAE don’t have access to PII and PHI • Right to audit by regulatory body (DoH). Clinical Use Overview: • Population / Intended User: Adult patients with liver tumors, including unresectable liver tumors, where focal liver-directed treatment is clinically appropriate. • To be Performed By: Trained physicians, typically interventional radiologists or other qualified liver tumor ablation specialists, supported by anesthesia and trained clinical staff. • Clinical Setting: Hospital-based interventional radiology suite, procedure room, or operating room with appropriate imaging, anesthesia, emergency, and infection-control support. • Condition of Use: 1. Patients with inadequate functional liver reserve. 2. Patients with Child-Pugh Class C liver disease. 3. Patients with decompensated liver cirrhosis. 4. Patients who cannot tolerate the planned anesthesia. 5. Tumors that are beyond the treatment depth or reach of the Edison System. 6. Inability to visualize the histotripsy bubble cloud for in vivo crosshair calibration. 7. Acoustic blockage or obstruction that prevents generation or maintenance of a sustained bubble cloud. 8. Treatment plans that exceed the system's thermal or safety safeguards. 9. Situations where the physician cannot safely establish or maintain the planned treatment volume (PTV) due to technical or anatomical limitations. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2026-07-22 | HTA26048 | Eight Sleep Pod - Sleep Apnea Diagnosis v1 | EIGHT SLEEP MIDDLE EAST LTD | Eight Sleep | United States of America | Medical Device | Diagnostics | DOH HTA decision: Conditional Approval • Conditional Approval under the DOH HTA framework for medical devices allows early market access in Abu Dhabi for highly innovative and/or clinically necessary medical devices, while ensuring appropriate safeguards for safety, effectiveness, and ongoing evidence generation where required. • Disclaimer: The Department of Health reserves the right to reassess, amend, suspend, or withdraw this conditional approval at any time in response to changes in UAE, FDA, or other relevant regulatory status, or emerging safety, effectiveness, or performance data, and may require additional HTA evaluation as deemed necessary. Key facts • Device Information: Eight Sleep Pod - Sleep Apnea Diagnosis v1 • Population / Intended Users: Adults with suspected OSA • To be Performed By: Self-administered by patients • Clinical Setting: Home environment • Approved Technology: Eight Sleep Pod, used with the manufacturer approved Eight Sleep Pod hardware, sensors, hub, software, and firmware configuration. • Regulatory status: o FDA: Applied for FDA clearance. o UAE regulatory status: Not registered with EDE The Eight Sleep Pod OSA functionality, when used with the manufacturer-approved Eight Sleep Pod hardware, sensors, hub, software, and firmware configuration, is to be used as a prescription (Rx) SaMD, upon physician order, to provide an Apnea-Hypopnea Index (AHI) score as an aid in the diagnosis of obstructive sleep apnea (OSA) in adults, subject to the conditions outlined below. Product-Specific Conditions 1. Intended Use: The SaMD is intended for adjunctive use only and is not a standalone diagnostic tool or a replacement for standard polysomnography. • All outputs must be reviewed and interpreted by a qualified healthcare professional. SaMD-generated results, including AHI estimates, shall be interpreted in conjunction with the patient’s clinical symptoms, medical history, physical examination, and clinician assessment. • In patients with suspected severe or complex sleep disorders, significant cardiopulmonary or neurological comorbidities, inconclusive findings, or discordance between clinical assessment and device results, confirmatory testing using the current standard of care (e.g. polysomnography or a validated home sleep apnea test) is required prior to clinical decision-making. • The Eight Sleep Pod shall be used only with manufacturer-approved hardware, sensors, hub, software versions, firmware updates, accessories, and consumables applicable to the approved OSA functionality. Unauthorized modifications or third-party components are not permitted. 2. AAMEN conditions to be fulfilled: For further details kindly contact @Abu Dhabi Healthcare Information Security Program: • AAMEN Exemption until 31st December 2026. • Condition: Right to audit by regulatory body (DOH). 3. Reassessment: Within one year of implementation/use in Abu Dhabi, DoH will conduct a formal reassessment of the conditional approval based on the evidence provided, to address remaining uncertainties and support system-level adoption. 4. Real-World Evidence Requirements: RWE submitted after one year of use shall include: • UAE healthcare setting: Evidence generated from or applicable to the UAE healthcare setting, including assessment of clinical performance, patient adherence, clinical impact, and integration into relevant clinical workflows, including usability and performance in relevant UAE patient subgroups. • Home-use validation: Evidence of performance under routine, unsupervised home conditions, including diagnostic concordance with polysomnography (PSG) or a validated home sleep apnea test (HSAT). • Physiological completeness: Evidence addressing the impact of the absence of blood oxygen saturation (SpO₂) measurement, including limitations in distinguishing apnea from hypopnea events and assessing hypoxic burden. • Broader clinical population applicability: Evidence in patients with conditions that may alter cardiorespiratory signals, including atrial fibrillation, heart failure, neuromuscular disease, chronic opioid use, and significant pulmonary disease. • Sample size and subgroup validation: Confirmation of performance in larger and more diverse populations, including relevant subgroups such as OSA severity, sex, and BMI. The validation sample shall include both device-positive and device-negative patients undergoing confirmatory testing to minimize verification bias and enable reliable estimation of false-positive and false-negative performance. • Usability-related real-world conditions: Evaluation of performance in multi-occupant beds, varying mattress types, presence of pets, and night-to-night variability under routine home use, including technically invalid studies and device/software failures. • Diagnostic performance: Assessment of sensitivity and specificity at AHI thresholds of ≥5, ≥15, and ≥30 events/hour, false-positive and false-negative rates, AHI agreement with PSG/validated HSAT, and OSA severity classification. Adequate validation at the AHI ≥5 events/hour threshold shall be provided to support the broader intended use as an aid in the diagnosis of OSA. • Algorithm robustness and generalizability: Assessment of real-world performance variability and the impact of dependence on annotator scoring patterns used in the training dataset. 5. RWE Acceptance Criteria & Submission: The submitted evidence should demonstrate clinically acceptable performance consistent with the evidence supporting this conditional approval, without clinically significant deterioration in diagnostic performance or newly identified safety concerns. The RWE submission should include the study methodology, population characteristics, performance outcomes, safety findings, and relevant statistical analyses. General Requirements (Standard DoH Conditions) 1. The technology shall comply with: • The conditions and regulations of the Department of Health regarding product marketing and advertising in the health sector. • All relevant laws, regulations, and standards of the Department of Health. • The Standard on Medical Device Reporting (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health, and a pharmacovigilance focal point shall be assigned for medical product reporting. 2. Physicians, healthcare professionals, or other personnel using the technology shall undergo appropriate operational or specialized training and obtain the required certification(s), whenever applicable. 3. A proper quality monitoring process shall be established, including the reporting of any adverse events, unwarranted consequences, or safety issues affecting patients or personnel. • Regular updates and reports regarding the product shall be provided to DoH upon request. • The device shall be installed, operated, and maintained in accordance with the manufacturer’s instructions, including environmental requirements, network connectivity, electrical safety, routine maintenance, and software updates. • Quality assurance processes shall include routine verification of sensor performance, software functionality, cybersecurity controls, cloud connectivity, and manufacturer-recommended calibration and maintenance procedures. DoH HTA approval does not constitute reimbursement, coverage, pricing, or procurement approval.
2026-06-10 | HTA26023 | Sinocare iCan i3 Continuous Glucose Monitoring System | Al mazroui medical and chemical supplies | Changsha Sinocare Inc | China | Medical Device | Others | Approved - Continuous Glucose Monitoring System is a real-time, continuous glucose monitoring device indicated for the measurement of glucose in interstitial fluids for people 2 years and older. It is intended to replace fingerstick blood glucose testing for diabetes treatment decisions. The System also detects trends and tracks patterns, and aids in the detection of episodes of hyperglycemia and hypoglycemia, facilitating both acute and long-term therapy adjustments. Interpretation of the System results should be based on the glucose trends and several sequential readings over time. The System can be used in conjunction with smart devices where the user manually controls actions for therapy decisions. Recommendation: • Use according to the approved indication and intended population (patients ≥2 years of age). • Ensure adequate patient and caregiver training on device use, alerts, and interpretation of glucose readings. • local post-market monitoring to assess safety, effectiveness, and user satisfaction in real-world practice. Clinical overview Population: measurement of glucose in interstitial fluids for people 2 years and older To be performed by: Healthcare Professionals/ Patients Clinical Setting: Healthcare facilities / Patients level Contraindication: Do not use the iCan i3 CGM if you are pregnant, on dialysis, implanted with a pacemakers or critically ill. It is not known how different conditions or medications common to these populations may affect performance of the system. iCan i3 CGM results may be inaccurate in these populations. No MRI/CT/Diathermy Do not wear your iCan i3 CGM (sensor, transmitter, receiver, or mobile device) for magnetic resonance imaging (MRI), computed tomography (CT) scan, or high-frequency electrical heat (diathermy) treatment. The iCan i3 CGM has not been tested in those situations. The magnetic fields and heat could damage the components of the iCan i3 CGM, which may cause it to display inaccurate sensor glucose results or may prevent alerts. Without iCan i3 CGM results or alert notifications, you might miss a severe low or high glucose event. To get the most out of your session, we advise that you try to schedule your procedure near the end of your sensor session to avoid needing an extra sensor. Please consult your doctor for recommendations with any other medical procedures. Bring your Blood Glucose Meter with you when you go to your procedure. Do not use iCan i3 CGM, if you are suffering from coagulation disorders or taking anticoagulant drugs. AAMEN Approved : 1. Health Information to be restricted. 2. L1 and L2 support shall be available within UAE. 3. Right to audit by regulatory body (DOH).
2026-06-10 | HTA26024 | Sinocare iCan i6 Continuous Glucose Monitoring System | Al mazroui medical and chemical supplies | Changsha Sinocare Inc | China | Medical Device | Others | Approved - Continuous Glucose Monitoring System (CGM System) is a real time, continuous glucose monitoring device indicated for the measurement of glucose in interstitial fluids for people 2 years and older. It is intended to replace fingerstick blood glucose testing for diabetes treatment decisions. The CGM System also detects trends and tracks patterns, assisting in the detection of episodes of hyperglycemia and hypoglycemia and facilitating both acute and long-term therapy adjustments. Interpretation of the CGM System results should be based on the glucose trends and several sequential readings over time. 5 The CGM System can be used in conjunction with smart devices, where the user manually controls actions for therapy decisions. Recommendation: • Use according to the approved indication and intended population (patients ≥2 years of age). • Ensure adequate patient and caregiver training on device use, alerts, and interpretation of glucose readings. • local post-market monitoring to assess safety, effectiveness, and user satisfaction in real-world practice. Clinical overview Population: measurement of glucose in interstitial fluids for people 2 years and older To be performed by: Healthcare Professionals/ Patients Clinical Setting: Healthcare facilities / Patients level Contraindication: Do not use the iCan i6/H6 Continuous Glucose Monitoring System if you are pregnant, on dialysis, implanted with a pacemaker, or critically ill. It is not yet clear whether the specific health conditions or commonly used medications of these individuals will impact the system's performance. In such cases, there may be errors in the monitoring results of iCan i6/H6. • No MRI/CT/high-frequency electrical heat (diathermy) Do not wear the iCan i6/H6 CGM (sensor, receiver, or mobile device) for magnetic resonance imaging (MRI), computed tomography (CT) scan, or high-frequency electrical heat (diathermy). The iCan i6/H6 CGM has not been tested in those situations. The magnetic fields and heat could damage the components of the iCan i6/H6 CGM, which may cause it to display inaccurate sensor glucose results or may prevent alerts. Without iCan i6/H6 CGM results or alert notifications, you might miss a severe low or high glucose event. To get the most out of your sensor, we advise that you try to schedule your procedure after the end of your sensor 8 session to avoid needing an extra sensor. Please consult your doctor for recommendations with any other medical procedures. Bring your blood glucose meter with you when you go to your procedure. Do not use iCan i6/H6 CGM, if you are suffering from coagulation disorders or taking anticoagulant drugs. AAMEN Approved: 1. Health Information to be restricted. 2. L1 and L2 support shall be available within UAE. 3. Right to audit by regulatory body (DOH).
2026-06-10 | HTA26032 | CyberKnife Treatment Delivery System - CyberKnife® S7 System | Al Naghi Medical | Accuray Inc | USA | Medical Device / System / Software-assisted radiation treatment delivery system | Therapeutic | Approved for adoption and for treatment planning and image-guided stereotactic radiosurgery and precision radiotherapy for tumors, lesions and conditions anywhere in the body when radiation treatment is indicated. The CyberKnife Systems may be used to treat astrocytoma, glioma, skull base tumors, metastases (brain and bony), nasopharyngeal carcinoma, meningioma, acoustic neuroma, schwannoma, pituitary adenoma, hemangioblastoma, craniopharyngioma, arteriovenous malformation, cavernous malformation, trigeminal neuralgia, and tumors of the neck, spine, pancreas, liver, lungs, ovary, prostate, and bladder. Patients should be examined by a team of physicians to determine if they are candidates for CyberKnife treatment. Recommendations 1. CyberKnife S7 is recommended for hospital use in radiation oncology departments for image-guided SRS/SBRT and precision radiotherapy where clinically indicated. 2. Use must be restricted to licensed healthcare facilities with appropriate radiotherapy shielding, radiation safety approval, and medical physics support. 3. Treatment must be prescribed and supervised by qualified radiation oncologists. 4. Treatment planning and verification must be performed by trained medical physicists/dosimetrists. 5. Operation must be performed by trained radiation therapists and competent CyberKnife users only, and in alignment and compliance with FANR regulations. 6. Each case should undergo multidisciplinary review, especially for complex tumors, re-irradiation, moving targets, and high-risk anatomical sites. 7. Patient-specific, plan verification, imaging verification, and dose constraints must be mandatory before treatment. 8. CyberKnife should be used according to approved indication/IFU and institutional SRS/SBRT protocols. 9. Clinical outcomes should be monitored, including local control, toxicity, treatment completion, retreatment, and adverse events. 10. A post-market surveillance and incident reporting process must be maintained through the UAE agent/PV focal point. All adverse events must be addressed to PVE@doh.gov.ae 11. CyberKnife use should comply with UAE ICT/data privacy requirements when patient data are processed or transferred. 12. Approved by AAMEN, the system logs must be accessible to the health entity for data compliance. For more information kindly contact AAMEN team via aamen@doh.gov.ae Clinical Use Overview Population / Intended User: Patients with tumors, lesions, or conditions requiring image-guided stereotactic radiosurgery or precision radiotherapy. To be Performed By: Radiation oncologist, medical physicist, dosimetrist, radiation therapist, with support from neurosurgery/oncology teams when required. Clinical Setting: Hospital-based radiation oncology department with shielded treatment vault and radiotherapy infrastructure as per FANR regulations. Condition of Use: Prescription-only, professional use only, patient-specific treatment planning, image-guided delivery, QA-verified treatment. To provide treatment planning and image-guided stereotactic radiosurgery and precision radiotherapy for tumors, lesions and conditions anywhere in the body when radiation treatment is indicated. The CyberKnife Systems may be used to treat astrocytoma, glioma, skull base tumors, metastases (brain and bony), nasopharyngeal carcinoma, meningioma, acoustic neuroma, schwannoma, pituitary adenoma, hemangioblastoma, craniopharyngioma, arteriovenous malformation, cavernous malformation, trigeminal neuralgia, and tumors of the neck, spine, pancreas, liver, lungs, ovary, prostate, and bladder. Patients should be examined by a team of physicians to determine if they are candidates for CyberKnife treatment. Contraindications: • Use should be avoided or reconsidered where radiation therapy is not clinically appropriate • Use should be avoided or reconsidered where prior radiation dose exceeds safe tolerance • Use should be avoided or reconsidered where target localization cannot be achieved reliably • Use should be avoided or reconsidered where patient condition prevents safe treatment positioning
2026-06-10 | HTA26033 | HEMORO ONE BLUE / HemoRO 4 ONE | ONE HEALTH LLC – DUBAI BRANCH | DWA GmbH & Co. KG, Großer Sand 8, 76698 Ubstadt-Weiher, Germany | Germany | Medical Device | Others | Approved to produce dialysis water by means of reverse osmosis and provide it to one or more dialysis equipment connected directly or through Media Supply Systems. Adoption should remain limited to this infrastructure/support role and not be interpreted as a standalone therapeutic or diagnostic indication. Recommendations: HemoRO 4 ONE / HemoRO 4 ONE BLUE is recommended for hospital and approved home-haemodialysis use as a single-patient reverse osmosis system for production and supply of dialysis water where such infrastructure is clinically required. It should be operated only by instructed and competent personnel, with manufacturer/service training as applicable. The technology should be used strictly according to IFU and only for the approved purpose of producing and supplying dialysis water. Routine microbiological and chemical monitoring of feed water and produced dialysis water should be mandatory, in line with ISO 23500-based facility practice. Validated disinfection procedures must be implemented consistently: heat disinfection for HemoRO 4 ONE where applicable chemical disinfection for HemoRO 4 ONE BLUE Preventive maintenance, annual electrical safety checks, and timely fault rectification should be incorporated into dialysis unit maintenance protocols. The device should be used only with approved accessories, parts, transformers, and cables, and connected dialysis machines must be disconnected during descaling / disinfection cycles as specified in the IFU. Clinical quality reporting should include: dialysis water quality compliance microbiological/endotoxin results device faults/alarms maintenance and disinfection records adverse events / field safety issues as part of routine post-market surveillance. Clinical Use Overview: Population / Intended User: Dialysis patients requiring haemodialysis support; intended operators are instructed healthcare professionals in hospitals/ICUs/dialysis centres, and in home environments the patient may also be the operator. To be Performed By: Trained dialysis staff / instructed operating personnel; maintenance, repair, and technical inspections only by authorized trained maintenance personnel. Clinical Setting: Dialysis centres, nephrology hospital units, intensive care units, and approved home haemodialysis settings. Condition of Use: Continuous-use accessory device for haemodialysis support, operated under the manufacturer’s specified installation, feed-water, maintenance, and monitoring requirements. Contraindications: No specific contraindications are stated because the unit is an accessory to dialysis machines. However, use outside intended purpose, with non-specified media, defective safety features, or improper installation/maintenance is contraindicated in practice.
2026-06-10 | HTA26034 | PATHFAST hs-cTnI | Al Shomoukh Trading for Technical & Medical Supplies – ASTEMED L.L.C | Pathfast | Japan | Diagnostic Test | Diagnostics | Approved for PATHFAST hs-cTnI is a product for in-vitro diagnostic use with the in vitro diagnostic (IVD) automated analyser PATHFAST for the quantitative measurement of cardiac Troponin I (cTnI) in human heparinized or EDTA whole blood and plasma. PATHFAST hs-cTnI is intended to be used: - as an aid in the diagnosis of acute coronary syndromes (ACS), - in the risk stratification in patients presenting with suspected ACS, - by laboratory technician, nurse or physician, - in hospital including emergency room, doctor’s office and clinical laboratory. PATHFAST hs-cTnI is a device for near patient testing (NPT). Recommendation: • Use only manufacturer-approved PATHFAST reagents, cartridges, calibrators, and consumables; reuse of cartridges or consumables is prohibited. • The device should be installed and operated in appropriate clinical environments according to the manufacturer’s instructions, including proper storage, calibration, and temperature conditions. • Results should always be interpreted together with patient symptoms, ECG findings, and clinical assessment, and not used as the sole basis for diagnosis. • Calibration and quality control should follow the manufacturer’s recommendations, including regular user calibration and routine QC testing with every new reagent lot. • Post-market monitoring should include documentation of invalid runs, turnaround time, discordant clinical results, maintenance issues, recalls, and user-reported malfunctions or adverse events. Clinical Use Overview: A rapid near-patient high-sensitivity cardiac troponin I assay used for early diagnosis and triage of patients with suspected acute coronary syndrome (ACS) and NSTEMI. Population / Intended User: Adults with suspected ACS/NSTEMI; intended users are trained healthcare professionals. To be Performed By: Laboratory technicians, nurses, physicians, or other trained healthcare professionals experienced in emergency or laboratory testing procedures. Clinical Setting: Healthcare fsacilities Condition of Use: PATHFAST hs-cTnI is a product for in-vitro diagnostic use with the in vitro diagnostic (IVD) automated analyser PATHFAST for the quantitative measurement of cardiac Troponin I (cTnI) in human heparinized or EDTA whole blood and plasma. PATHFAST hs-cTnI is intended to be used: - as an aid in the diagnosis of acute coronary syndromes (ACS), - in the risk stratification in patients presenting with suspected ACS, - by laboratory technician, nurse or physician, - in hospital including emergency room, doctor’s office and clinical laboratory. PATHFAST hs-cTnI is a device for near patient testing (NPT). Contraindications: No explicit patient contraindications were identified. Main restrictions relate to operational use, including use only with approved consumables and reagents, proper sample handling, and adherence to manufacturer instructions and laboratory quality procedures.
2026-06-10 | HTA26035 | Transcutaneous Jaundice Detector - MBJ20 | Hessa Medical Equipments Co LLC | Beijing M&B Electronic Instruments Co., Ltd | China | Medical Device | Diagnostics | Approved for adoption use as examination of neonate jaundice, by determining the concentration of bilirubin transcutaneously. It is used for the initial screening and monitoring of bilirubin values in jaundice treatment and not to be used as the sole diagnostic tool Recommendations: MBJ20 Transcutaneous Jaundice Detector is recommended for hospital use for neonatal jaundice screening and monitoring where bilirubin assessment is clinically indicated. To be operated by trained and competent healthcare professionals only. Use only in neonates, preferably on forehead or sternum. Serum bilirubin confirmation is required for babies <24 hours old, <35 weeks gestation, suspected hemolysis/pathological jaundice, TcB near treatment threshold, or before phototherapy/escalation decisions. Do not use during MRI or near strong electromagnetic interference. Avoid direct light emission into eyes. During/after phototherapy, use protected non-exposed measurement sites; phototherapy lights should be turned off during measurement. Implement cleaning, calibration, battery replacement, routine inspection every 3 months, and annual preventive maintenance. Include device use in NICU/newborn jaundice screening protocols. Monitor false high/low readings, serum confirmation rates, phototherapy escalation, and adverse/device malfunction events as part of PMS and quality reporting. Clinical Use Overview: Population / Intended User: Neonates requiring jaundice screening or monitoring. To be Performed By: Doctors, nurses, neonatal/pediatric clinical staff trained in TcB use. Clinical Setting: Hospitals, NICU, newborn nursery, obstetric/neonatal/pediatric departments. Condition of Use: Examination of neonate jaundice, by determining the concentration of bilirubin transcutaneously. It is used for the initial screening and monitoring of bilirubin values in jaundice treatment. Contraindications / Exclusions: Not to be used as the sole diagnostic tool Not for use in adults or older children Avoid use during MRI procedures Avoid direct eye exposure Use with caution during/after phototherapy Use with caution in early hemolytic or pathological jaundice Limited evaluation under high bilirubin (>20 mg/dL) conditions Do not use on neonates with skin conditions, bruising, edema, or lesions at the measurement site. Contraindicated in cases where transcutaneous bilirubin measurement may provide unreliable results, including severe skin pigmentation abnormalities or significant dermal discoloration. The device should not replace serum bilirubin testing when urgent or confirmatory diagnosis is clinically required. Do not use outside the manufacturer’s intended neonatal population and validated clinical settings. Contraindicated if the device has not undergone proper calibration or maintenance procedures.
2026-04-28 | HTA25109 | MindMotion™ GO (MMGO) | Amana Healthcare, Abu Dhabi | MindMaze SA, Chemin de Roseneck 5, Lausanne, Switzerland | Switzerland | System/ Software/ AI / App | Therapeutic | Conditonally approved for a software used with motion tracking solutions to support the physical rehabilitation and motor recovery of adults with chronic stroke or Parkinson’s disease. The device is intended for use under the direct or remote supervision of a healthcare professional, either in clinical settings or at a patient’s home. Ensure the device is used only under supervision of a healthcare professional, especially for home-based rehabilitation. Monitor and document clinical outcomes and safety events (e.g., falls, dizziness, fatigue) as part of institutional quality assurance. Use the system as an adjunct to standard rehabilitation care, not a replacement, particularly to increase therapy intensity and patient engagement through gamified exercises. Condtionally Approved as per AAMEN : - PHI/PII shall not be transferred or accessible outside the borders of UAE - MWA approval is required - Right to Audit by regulatory (DoH) Population: Minimum age: 18 years old. Prevalent physical condition: • Motor impairments of the hand, upper-extremity, trunk and/or lower extremity due to stroke in the chronic phase; or • Lower-limb impairments related to gait, balance, and/or trunk stability due to Parkinson’s disease. • The patient must be medically stable, and they must be able to actively participate in the therapy delivery (i.e. understand and follow instructions, and perform training defined by the healthcare professional). Minimum knowledge required: Understand instructions provided by the healthcare professional To be performed by: • Minimum education level: Physiotherapist, occupational therapist, physiatrist, neuropsychologist or similar healthcare qualification. • Professional competency: The healthcare professionals should be able to setup the MindMotion GO in a suitable environment, choose the correct training program to be followed for a specific patient, and monitor the patient training activity. For such activities, the healthcare professionals use all the components of the MindMotion GO, including the associated hardware peripherals, and the Companion web application. • In case of use of MindMotion GO at home, a third-party care giver or support person (e.g., family member) can provide help to set up MindMotion GO and in case of technical problems. • Minimum age: 18 years old. • Minimum knowledge required: Basic computer literacy, ability to understand instructions provided by the healthcare professional. Clinical setting: Healthcare facilities, at home. Condition of use: MindMotion GO is intended for patients with: • motor impairment of the hand, the upper extremity, trunk, and/or lower extremity in chronic Stroke, or • lower-limb impairments related to gait, balance, and/or trunk stability in Parkinson’s disease.
2026-04-28 | HTA26025 | Fully Automated Molecular POCT Analyzer - HysenDx 200 / DX200 | Omics One Sciences Medical Supplies Trading LLC | Hysen Biotech.Inc | Republic of Korea / Korea (South) | Diagnostic Test | Diagnostics | Approved for adoption and used as quantitative detection of analytes in the related pathogen nucleic acids (DNA/RNA) from human samples using the polymerase chain reaction process. The samples can be nasopharyngeal and oropharyngeal samples, whole blood, blood plasma, blood serum, saliva samples and so on. The instrument is for in vitro diagnostic only. The Fully Automated Molecular POCT Analyzer is intended for used in medical and biological laboratories by professional user trained in molecular biological techniques and the operation of Fully Automated Molecular POCT Analyzer. Recommendations: 1. Hysen Fully Automated Molecular POCT Analyzer (HysenDx 200) is recommended for restricted hospital laboratory use for rapid molecular detection of pathogen nucleic acids, in line with the IFU and validated compatible assays only. 2. To be operated only by trained and competent healthcare laboratory professionals experienced in molecular testing and biosafety procedures. 3. Use only manufacturer-provided or manufacturer-recommended consumables, chips, and accessories; reuse of chips/consumables is prohibited. 4. The device should be installed and used only in appropriate indoor laboratory clinical environments with proper grounding, ventilation, spacing, and temperature/humidity controls. 5. Operators must observe all biohazard and hot-surface precautions, including PPE, avoidance of touching heated modules, and adherence to PCR laboratory regulations. 6. The analyzer should be integrated into infectious disease - respiratory testing pathways where rapid molecular results can influence isolation, cohorting, treatment selection, and antimicrobial stewardship. 7. Initial installation, annual recalibration, maintenance, and troubleshooting should be supported by the manufacturer or authorized service personnel only. 8. Post-market monitoring should include documentation of invalid-failed runs, turnaround time, concordance with central laboratory testing, contamination incidents, maintenance issues, and user-reported adverse events/malfunctions. Clinical Use Overview: A rapid, near-patient or decentralized molecular PCR analyzer for infectious disease diagnosis, especially respiratory pathogen detection when a fast sample-to-result workflow is needed. Population / Intended User: Patients suspected of infectious disease; end users are trained laboratory/healthcare professionals. To be Performed By: Trained professional users in medical and biological laboratories; practically this includes laboratory personnel and other trained healthcare professionals according to institutional policy. Clinical Setting: Hospitals, medical laboratories, biological laboratories, and suitable decentralized healthcare settings where molecular POCT is operationally appropriate. Condition of Use: The Fully Automated Molecular POCT Analyzer is an automated instrument used for quantitative detection of analytes in the related pathogen nucleic acids (DNA/RNA) from human samples using the polymerase chain reaction process. The samples can be nasopharyngeal and oropharyngeal samples, whole blood, blood plasma, blood serum, saliva samples and so on. The instrument is for in vitro diagnostic only. The Fully Automated Molecular POCT Analyzer is intended for used in medical and biological laboratories by professional user trained in molecular biological techniques and the operation of Fully Automated Molecular POCT Analyzer. Contraindications: No explicit patient contraindications. The main restrictions are operational: not for use outside intended purpose, not with non-approved consumables, not in hazardous environments, and not with prohibited flammable/explosive/reactive materials in the centrifuge component.
2026-04-28 | HTA26028 | gekoTM T-3 Neuromuscular Stimulator | Medlab Medical Equipment Trading LLC | Firstkind Limited | United Kingdom | Medical Device | Therapeutic | Approved for Increasing local blood circulation, Edema reduction, Immediate post-surgical stimulation of the calf muscles to prevent venous thrombosis and Stimulation of the calf muscles to prevent venous thrombosis is non-surgical patients at risk for venous thromboembolism. • Use as an alternative if patients cannot tolerate compression devices. • Ensure proper training on device placement and use. Population: Non-surgical patients at risk for venous thromboembolism To be performed by: Healthcare professional. Clinical setting: Healthcare facilities or home Condition of use: It is to be used on a single patient, in hospital, clinic and home environments. Apply to one or both legs as advised by your Healthcare Professional (e.g. Nurse or Doctor). Patients should be trained in the positioning and operation of the device. The geko device is intended for up to 24 hours continuous use, until it is no longer needed. It may be removed for short periods of time when washing or having various tests (see warnings and precautions). Each device is single use and should be replaced every 24 hours, the device has an additional 6 hours run time.
2026-03-26 | HTA26022 | MESI m-Tablet System | SKY SMARTMED MEDICAL SURGICAL ARTICLES & REQUISITES TRADING L.L.C | MESI Medical | Sloveinia | Medical Device | Diagnostics | Approved for use on adult subjects at risk of having or developing peripheral arterial disease (PAD). MESI mTABLET ABI is intended for the rapid measurement of ankle-brachial pressure index (ABI) and pulse volume plethysmography in adults. It is suitable for use in wound care assessment, for assessing symptomatic PAD, and as a screening device for PAD. It may also be used on patients with venous or arterial ulcers prior to application of compression therapy. MESI mTABLET ABI can be used on patients with unilateral lower limb amputation. The MESI mTABLET ABI is intended to be used to spot-check patients. The MESI mTABLET ABI provides information regarding patient risk. The physician has the responsibility of making proper judgments based on this information. • Use the MESI mTABLET device for peripheral arterial disease (PAD) screening • Use the device as a rapid ABI measurement tool when Doppler assessment is not readily available. • Use the device as an initial screening tool, while confirming abnormal results with standard Doppler or further vascular imaging if needed. • Use the device to support early detection of peripheral arterial disease (PAD), which may improve timely diagnosis and management of vascular disease. Approved as per AAMEN: - Support (L1, L2 and L3) shall have no visibility of PHI/PII. - L1 and L2 support shall be within UAE. - Right to Audit by Regulatory (DoH) Clinical Overview: Population: use on adult subjects at risk of having or developing peripheral arterial disease (PAD) Clinical setting: Healthcare facilities
2026-03-26 | HTA26021 | Pteye Parathyroid Detection System | Pioneer medical Supplies LLC | Medtronic | USA | Medical Device | Diagnostics | Approved for : The PTeye™ system is an adjunctive tool intended to aid in the identification of parathyroid tissue by confirming parathyroid tissue already visually located by the surgeon in adults 18 years of age and older. Intended use The PTeye™ system is for prescription use. The intended end-user of the PTeye™ device is any surgeon performing thyroid or parathyroid procedures who desires an aid for parathyroid identification. The intended environment of use is any OR theater where thyroid or parathyroid procedures are being performed. • Calibrate before use • Use steady contact • Minimize interference • Confirm with clinical judgment Population: Adult patients undergoing thyroidectomy or parathyroidectomy procedures. Clinical Setting: Healthcare facilities Condition of use :As an adjunctive tool to confirm parathyroid tissue that the surgeon has already visually identified (or strongly suspects) during the procedure. It provides real-time feedback on tissue fluorescence. Contraindications: Not intended for pediatric use.
2026-03-26 | HTA26017 | URO17 Bladder Cancer Test - KDx-400 | NMC Group Services Ltd | KDx Diagnostics | United States of America | Diagnostic Test | Diagnostics | Conditionally approved for adoption and as adjunct tool that utilizes immunostaining with a chromogenically labeled monoclonal antibody for the qualitative determination of Keratin 17 in urine cytology specimens from patients with bladder cancer or hematuria. Results of the URO17® Bladder Cancer Test are intended for use as a noninvasive method for monitoring for tumor recurrence or triaging patients into cystoscopy in conjunction with other clinical diagnostic methods (e.g., cystoscopy, cytology). Recommendations 1. URO17® is recommended as an adjunctive non-invasive urine biomarker test for detection of bladder cancer in patients with haematuria or lower urinary tract symptoms suspicious for malignancy, and for surveillance of recurrence in urothelial cancer follow-up. 2. The test should be requested and interpreted only by trained urologists, oncologists, and/or pathologists within an appropriate clinical pathway. 3. URO17 results should not be used as a stand-alone replacement for cystoscopy, biopsy, or histopathology when these are clinically indicated. 4. Positive results should prompt appropriate urologic evaluation and correlation with cystoscopy, imaging, cytology, and histopathology as needed. 5. Negative results may be clinically helpful, especially given the high reported negative predictive value in early studies but should still be interpreted in the context of symptoms, risk factors, and clinician judgment. 6. Caution is advised in patients with prior intravesical therapy, inflammatory changes, atypia, prior BCG therapy, or radical radiotherapy, because diagnostic utility may be less certain in these settings. 7. Local implementation should include audit of diagnostic yield, false positive/false negative rates, downstream cystoscopy use, and patient outcomes. 8.Company to continuously monitor patients outcomes to ensure technology’s effectiveness and safety and report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division in DOH as part of a Real World Data. Clinical Use Overview Population / Intended User: Adults with haematuria, lower urinary tract symptoms suspicious for malignancy, or patients under surveillance for urothelial/bladder cancer recurrence. To be Performed By: Qualified laboratory personnel/pathology services, with results interpreted by urologists and/or pathologists. This is inferred from the immunocytochemistry laboratory workflow and intended clinical use. Clinical Setting: Hospital laboratories, reference laboratories, urology clinics, and oncology follow-up settings. Condition of Use: The URO17® Bladder Cancer test is a reagent set (antibody plus control slides) that utilizes immunostaining with a chromogenically labeled monoclonal antibody for the qualitative determination of Keratin 17 in urine cytology specimens from patients with bladder cancer or hematuria. Results of the URO17® Bladder Cancer Test are intended for use as a noninvasive method for monitoring for tumor recurrence or triaging patients into cystoscopy in conjunction with other clinical diagnostic methods (e.g., cystoscopy, cytology). Contraindications: The main limitation is that URO17 should not be used as the sole determinant to exclude or confirm malignancy where standard diagnostic procedures are indicated. Evidence is also less conclusive in patients post-BCG or post-radical radiotherapy.
2026-03-26 | HTA26016 | MONARCH PLATFORM | Johnson and Johnson Middle east | Auris Health, Inc. | USA | Medical Device | Therapeutic | Approved for : The MONARCH™ Bronchoscope and the MONARCH™ Platform and its accessories are intended to provide bronchoscopic visualization of and access to patient airways for diagnostic and therapeutic procedures. Recommendation: • Use as an advanced diagnostic tool for patients with suspected lung lesions/nodules where tissue sampling is needed. • Combine with appropriate adjunct imaging/techniques when available. • Implement quality monitoring Approve as per AAMEN team: - PHI/PII shall not leave UAE. - Support (L1, L2 and L3) shall have no visibility of PHI/PII. - L1 and L2 support shall be within UAE Clinical Overview: Population: Adult patients require diagnostic evaluation of suspected or confirmed lung lesions. Performed by: Physician operator/ healthcare professionals Clinical Setting: healthcare facilities Condition of Use: The monarch platform and its accessories are intended to provide bronchoscopic visualization of and access to patient airways for diagnostic and therapeutic procedures. Contraindications: Contraindications include but are not limited to: • Patient whose general health or respiratory function or both are compromised to the point that the patient would not tolerate flexible bronchoscopy. • Absence of a trained bronchoscopist to perform or closely and directly supervise the procedure, as well as manage complications common to flexible bronchoscopy. • Use of the System in patients with electrically or magnetically activated implanted medical devices.
2026-03-26 | HTA26015 | Toda Amniodiag 5 Strip | CAMBRIDGE MEDICAL SURGICAL EQUIPMENT AND INSTRUMENTS TRADING LLC | TODA PHARMA | France | Medical Product/ Kit | Diagnostics | Toda Amniodiag 5 Strip is approved for a rapid test intended to detect the Insulin-like growth factor-binding protein 1 (IGFBP-1) antigen from clinical specimens. It’s detection in cervical secretions of pregnant women predicts an increased risk of preterm delivery. • Toda Amniodiag® 5 Strip is recommended for use as a rapid adjunctive diagnostic test to support the assessment of suspected premature rupture of membranes (PROM) when clinically indicated. • To be used by trained and competent healthcare professionals only • Sample collection should be performed prior to vaginal examination or procedures to reduce the risk of contamination and false results. • Caution is advised in cases of vaginal bleeding or contamination, as results may be affected; repeat or alternative testing may be necessary. • Storage, handling, timing of testing, and result interpretation must strictly follow manufacturer instructions to ensure accuracy. • Quality control, staff training, and ongoing monitoring of test performance, including invalid or discordant results, should be routinely implemented and documented. Population: Pregnant women (before 37 weeks of amenorrhea). To be performed by: This product is intended for professional use in laboratory and Point of Care environments, or self-collection under the supervision of a healthcare worker. The test is for professional use only and not for self-testing. Clinical setting: Healthcare facilities. Condition of use: Detection of IGFBP-1 during pregnancy
2026-03-26 | HTA26014 | Toda fFN Diag | CAMBRIDGE MEDICAL SURGICAL EQUIPMENT AND INSTRUMENTS TRADING LLC | TODA PHARMA | France | Medical Product/ Kit | Diagnostics | Approved for adoption and as a rapid adjunctive diagnostic tool, it is a simple and rapid immunochromatographic test for the qualitative detection of fetal fibronectin in cervicovaginal secretions. This test can be performed between the 22nd and 35th week of amenorrhea. The increase in the concentration of fetal fibronectin in cervicovaginal secretions during this period constitutes a marker of preterm delivery risk. Recommendations: 1. Toda fFN Diag® is recommended for hospital use as a rapid adjunctive diagnostic test to assess threatened preterm delivery risk in pregnant women when clinically indicated. 2. To be used by trained and competent healthcare professionals only. 3. The test should be integrated into obstetric triage, preterm labor assessment protocols and interpreted together with clinical findings. 4. Sample collection should occur before digital examination or cervical manipulation to reduce false-positive results. 5. The test should not be used as a standalone decision tool; if clinical symptoms persist despite a negative result, additional investigations are recommended. 6. Use restrictions, particularly in women with premature rupture of membranes, cervical dilatation ≥3 cm, suspected placental abruption, placenta previa, or moderate/severe vaginal bleeding. 7. Storage, handling, expiry, and infection-control precautions must be strictly followed. 8. Clinical outcomes and post-market performance, including admission avoidance, need for escalation, discordant results, and any incidents/complaints, should be routinely monitored and documented. Clinical Use Overview: Population / Intended User: Pregnant women, particularly between the 22nd and 35th week of amenorrhea; intended users are health professionals. To be Performed By: Healthcare professionals / lab or clinical staff trained in sample collection and test performance. Clinical Setting: Hospital obstetric / gynecology / emergency / maternity settings. The clinical test report notes suitability for large, medium, and small hospitals. Condition of Use: Toda fFN Diag® is a simple and rapid immunochromatographic test for the qualitative detection of fetal fibronectin in cervicovaginal secretions. This test can be performed between the 22nd and 35th week of amenorrhea. The increase in the concentration of fetal fibronectin in cervicovaginal secretions during this period constitutes a marker of preterm delivery risk. Contraindications: The test is not recommended in: • premature rupture of membranes • cervical dilatation ≥ 3 cm • suspected or diagnosed placental abruption • placenta previa • moderate or severe vaginal bleeding
2026-03-26 | HTA25117 | PRAMS and T-Bot - The Fertility Support Software | FertiClinic | Blue Door Fertility Limited (in association with ARGC) | United Kingdom | System/ Software/ AI / App | Diagnostic & therapeutic support software | Conditionally Approved for The software is intended for use by qualified fertility consultants and authorised clinical staff within fertility clinics. It is designed to capture, process, and present patient data, including laboratory test results and scan findings, and to generate data-driven advisory outputs such as suggested dose adjustments and monitoring schedules. The software functions solely as a clinical decision support tool to aid, but not replace, professional medical judgement. It is not intended to make autonomous treatment decisions or to directly administer therapy. Recommendation: Use only by trained staff: Ensure that only qualified fertility doctors and authorized clinic personnel use the software (as intended). Data entry accuracy: Enter all lab and ultrasound results fully and carefully every day. Double-check that values are correct before T Bot runs. Review suggestions: Always read T Bot’s advice in full. Cross-check any proposed medication changes or timings with your clinical judgment. Do not follow a suggestion blindly. Verify orders: Before issuing any prescription or patient instructions, cross-check that the instructions match intended doses. Follow protocols: Continue to follow standard clinic protocols for monitoring and patient communication, especially if any safety alerts appear. Technical safeguards: Use a secure clinic device (desktop, laptop, or tablet) with up-to-date antivirus and browser (Chrome or Edge). Do not share logins and enable automatic screen lock when idle. Keep software updated: Accept regular software updates (these are done centrally with release notes) so you have the latest safety fixes and algorithm improvements. Conditionally approved as per AAMEN team: 1. Obtain mobile and web application approval (Circular 68/2022) 2. Access to PHI data shall be restricted to Clinical Professionals only. 3. Right to audit by regulatory body (DOH). Clinical Overview: Population: The primary users of the software are fertility specialists (qualified medical doctors). The software may also be used by authorized healthcare professionals, including nurses and embryologists, for the purpose of entering, reviewing, and verifying patient data within the fertility treatment workflow. All diagnostic and therapeutic decision-making and treatment responsibility remain with the attending fertility specialist. Users: Qualified fertility specialists Clinical setting: The software is intended for use within fertility clinic environments under appropriate clinical supervision for patients. Validation has been performed within a clinical setting involving patients undergoing controlled ovarian stimulation as part of assisted reproductive treatments. Condition of use: Used during the ovarian stimulation phase of IVF cycles to guide dosing and timing decisions. It is only for patients actively undergoing IVF; it does not apply outside an IVF cycle.
2026-03-26 | HTA25102 | OMOM RC Robotic Capsule Endoscopy System - RC100 | Pioneer Medical Supplies LLC | Chongqing Jinshan Science & Technology (Group) Co., Ltd. | China | Medical Device | Diagnostic / Gastroenterology | Approved for adoption and for collecting and viewing images of stomach and small bowel and is used in medical institutions. It can be used for health management (examination), preliminary screening of gastric cancer, examination of small bowel mucosal damages and related diseases, especially for the following diseases: 1. Those who need gastroscopy but are unwilling to accept or cannot tolerate gastroscopy (including sedated gastroscopy); 2. Gastric examination of health management (physical examination) population. 3. Primary screening of gastric cancer. 4. Examination of gastrointestinal mucosal damage related to drugs (such as antiplatelet drugs, nonsteroidal anti-inflammatory drugs, etc.). 5. The reexamination or follow-up of some gastric lesions, such as varicose fundus, atrophic gastritis, gastric ulcer after standard treatment, gastric polyp, etc. 6. Follow-up of patients after partial gastrectomy and endoscopic minimally invasive treatment. 7. Observation and monitoring of Crohn's disease that cannot be detected by gastrointestinal endoscopy or colonoscopy. 8. Observation and monitoring of the lesions leading to overt occult GI bleeding, which cannot be detected by gastrointestinal endoscopy or colonoscopy. 9. Observation and monitoring of the lesions that may lead to iron deficiency anemia, which cannot be detected by gastrointestinal endoscopy or colonoscopy. Recommendations: 1. OMOM RC Robotic Capsule Endoscopy System is recommended for hospital use as a noninvasive diagnostic option for stomach and small bowel imaging in appropriately selected adult patients where capsule endoscopy is clinically indicated. 2. To be operated only by trained and competent healthcare professionals familiar with capsule endoscopy workflow, robotic console operation, patient selection, and image interpretation. 3. The technology should be used in licensed healthcare facilities with appropriate gastroenterology/endoscopy support and capability for escalation if capsule retention or other complications are suspected. 4. Strict patient selection and screening for contraindications must be performed before use, especially for obstruction, stenosis, fistula, swallowing difficulty, implanted electronic devices, pregnancy, and inability to undergo abdominal surgery if required. 5. The device should be positioned as a diagnostic complement/alternative and not a replacement for conventional gastroscopy when biopsy, tissue sampling, or therapeutic intervention is expected. 6. Bowel and gastric preparation protocols, hydration instructions, and examination workflow should be standardized locally before implementation. 7. Initial and periodic staff training and competency validation should be ensured for robotic movement unit use, recorder pairing, real-time monitoring, data download, and reporting through VUE Smart. 8. Clinical outcomes and safety events, including completion rate, lesion detection yield, repeat examination rate, capsule passage/retention, adverse events, and downstream need for conventional endoscopy, should be monitored as part of routine post-market surveillance and quality reporting. Clinical Use Overview: Population / Intended User: Adult patients requiring collection and viewing of stomach and small bowel images in medical institutions; especially those unwilling or unable to tolerate conventional gastroscopy. To be Performed By: Qualified medical personnel / trained physicians and healthcare professionals with specific technical training in system operation and capsule endoscopy interpretation. Clinical Setting: Hospitals and medical institutions, particularly gastroenterology/endoscopy services. Condition of Use: Prescription, professional-use diagnostic medical device; restricted to licensed healthcare professionals per UAE classification. Utilized for collecting and viewing images of stomach and small bowel, and is used in medical institutions. It can be used for health management (examination), preliminary screening of gastric cancer, examination of small bowel mucosal damages and related diseases, especially for the following diseases: 1. Those who need gastroscopy but are unwilling to accept or cannot tolerate gastroscopy (including sedated gastroscopy); 2. Gastric examination of health management (physical examination) population. 3. Primary screening of gastric cancer. 4. Examination of gastrointestinal mucosal damage related to drugs (such as antiplatelet drugs, nonsteroidal anti-inflammatory drugs, etc.). 5. The reexamination or follow-up of some gastric lesions, such as varicose fundus, atrophic gastritis, gastric ulcer after standard treatment, gastric polyp, etc. 6. Follow-up of patients after partial gastrectomy and endoscopic minimally invasive treatment. 7. Observation and monitoring of Crohn's disease that cannot be detected by gastrointestinal endoscopy or colonoscopy. 8. Observation and monitoring of the lesions leading to overt occult GI bleeding, which cannot be detected by gastrointestinal endoscopy or colonoscopy. 9. Observation and monitoring of the lesions that may lead to iron deficiency anemia, which cannot be detected by gastrointestinal endoscopy or colonoscopy. Contraindications: • No condition for surgery / refusal of abdominal surgery • Pacemaker unless MRI-compatible • Cochlear implant, magnetic metal drug infusion pump, nerve stimulator, other implanted electronic devices, magnetic metal foreign body • Pregnancy • Relative: GI obstruction, stenosis, fistula, dysphagia, mental disorder, under 18 or over 70 years, pathological disease states
2026-03-26 | HTA25101 | Velmeni | Smart Medical Solutions | Velmeni Inc. | United States | System/ Software/ AI / App | Image analyzer | Contionally approved for :VELMENI for DENTISTS (V4D) is a concurrent-read, computer-assisted detection software intended to assist dentists in the clinical detection of dental caries, fillings/restorations, fixed prostheses, and implants in digital bitewing, periapical, and panoramic radiographs of permanent teeth in patients 15 years of age or older. This device provides additional information for dentists in examining radiographs of patients’ teeth. This device is not intended as a replacement for a complete examination by the dentist or their clinical judgment that considers other relevant information from the image, patient history, or actual in vivo clinical assessment. Final diagnoses and patient treatment plans are the responsibility of the dentist. Recommendations: Adoption • Consider adopting V4D in dental centers to enhance radiographic detection, particularly where variability in diagnostic accuracy is a challenge. • Implement V4D as a decision-support tool, ensuring it complements—not replaces—clinical assessment. Verification & Implementation • Conduct local verification using facility radiographs to confirm acceptable performance before full deployment. Quality, Safety, and Monitoring • Monitor performance indicators: false positives/negatives, reader feedback, and detection consistency. Condtionally Approved as per AAMEN team: - Anonymization of PHI/PII. - Support (L1, L2 and L3) shall have no visibility of PHI/PII. - L1 and L2 support to be available within UAE. - Obtain DOH approval for Mobil/Web application. - Right to audit by regulatory body (DOH). Clinical Overview: Population: Patients ≥15 years with permanent dentition Performed by: Licensed dental professionals Setting: Dental clinics Condition of Use: not intended as a replacement for a complete examination by the dentist or their clinical judgment that considers other relevant information from the image, patient history, or actual in vivo clinical assessment. Final diagnoses and patient treatment plans are the responsibility of the dentist.
2026-02-26 | HTA26011 | Easy Puff T-Piece Infant Resuscitator | HESSA Medical Equipment’s Co. LLC | NOVOS Medical Systems | Türkiye | Medical Device | Therapeutic | Approved for Adoption Approved for adoption for use Indications are as follows: • Hypoxemia / hypoxia • Low apgar score and poor breathing effort • Heart rate lower than 100 per minute Recommendations 1. EasyPuff T-Piece Infant Resuscitator is recommended for hospital use in neonatal resuscitation where controlled positive pressure ventilation is clinically indicated 2. To be operated by trained and competent healthcare professionals only 3. Pre-set PIP and PEEP values must be verified before each use 4. Continuous monitoring of neonatal respiratory status must be ensured 5. Device functionality must be checked prior to every resuscitation 6. To be included in delivery room, NICU, and neonatal stabilization protocols 7. Initial and periodic staff training and competency validation should be ensured 8. Clinical outcomes and safety events should be monitored and documented as part of routine post-market surveillance Clinical Use Overview Population / Intended User EASYPUFF is intended to be used for neonates who need resuscitation. To be Performed By 1. Neonatologists 2. Pediatricians 3. Trained nurses 4. Respiratory therapists Clinical Setting Delivery rooms Neonatal Intensive Care Units Newborn nurseries Condition of Use Hospital-based use with pressurized gas source Contraindications Contraindications • No known contraindications when used according to the manufacturer’s Instructions for Use Restrictions for Use • Intended for use in infants with a maximum body weight of 10 kg • To be connected only to flow-regulated medical air, oxygen, or air-oxygen mixture • Input gas flow rate must be within 5–15 L/min (recommended 8 L/min) • Maximum pressure relief setting must not exceed 80 cmH₂O • Adjustment of maximum pressure relief valve only by trained neonatal resuscitation personnel • Use only with NOVOS-approved original accessories and components • Not intended for use with third-party or non-validated accessories +[@[Recommendations (for US Approval Sheet and Health Technology Review Report) Note: In case of Technology Approval for Market Entry, but w/ below SPECIFIC Conditions]]
2026-02-26 | HTA26010 | Novos Bilisphere 360 LED (Neonatal Phototherapy Device) | HESSA Medical Equipments Co LLC (UAE Agent/Company) | NOVOS MEDICAL SYSTEMS | Türkiye | Medical Device | Therapeutic | Approved for adoption and for the following; - Hyperbilirubinemia treatment - Kernicterus prevention - Patients with light sensitivity disorders (e.g., congenital erythropoietic porphyria) Recommendations 1. Novos Bilisphere 360 LED Phototherapy System is recommended for hospital use in the treatment of neonatal hyperbilirubinemia where phototherapy is clinically indicated. 2. To be operated by trained and competent healthcare professionals only, specifically neonatal care unit personnel trained on device operation and risks. 3. Eye protection must be applied throughout phototherapy in accordance with neonatal phototherapy safety practice and facility protocol. 4. Continuous neonatal monitoring must be ensured, including temperature and overall clinical status, with alarm limits used as applicable (air and skin temperature monitoring features are present in the device control interface). 5. Dose delivery must be optimized by ensuring appropriate irradiance and maximal exposed skin surface coverage, consistent with the device design intent to maximize both parameters. 6. To be included in NICU and newborn unit phototherapy protocols, aligned with clinical thresholds and escalation pathways (phototherapy as first-line management). 7. Initial and periodic staff training and competency validation should be ensured, including infection control and safe handling practices. (Professional-use requirement and risk management considerations.) 8. Clinical outcomes and safety events should be monitored and documented, including TSB reduction rate, duration of therapy, need for escalation, thermal instability events, skin events, and suspected device-related incidents, with reporting to the DOH PV program as required. Clinical Use Overview: Neonatal phototherapy device used in newborn services to reduce bilirubin levels in infants with neonatal jaundice/hyperbilirubinemia, providing blue light therapy with high coverage and irradiance. Population / Intended User: • Population: Term, preterm, post-term infants with hyperbilirubinemia. • Intended user: Trained medical personnel in neonatal care units (professional use only). To be Performed By: Neonatal nurses, neonatology clinicians, and trained NICU/newborn unit staff competent in phototherapy delivery and monitoring. Clinical Setting: NICU, newborn nursery, hospital neonatal units. Condition of Use: Use within facility phototherapy protocol with continuous monitoring and adherence to infection prevention and device safety measures (as per risk management and professional-use requirement). Contraindications: Contraindications • Phototherapy should be used with caution and based on clinical judgment in neonates with: o Conditions associated with photosensitivity o Rare metabolic or dermatological disorders where light exposure may be contraindicated • Use outside the intended neonatal population or clinical setting is not recommended. Potential Adverse Events The following adverse events are known and expected risks associated with neonatal phototherapy and are not unique to this device: Thermal Effects • Hyperthermia • Hypothermia Skin-Related Effects • Transient skin rash • Erythema • Skin dryness Fluid and Metabolic Effects • Increased insensible water loss • Dehydration if fluid intake is inadequate Gastrointestinal Effects • Increased stool frequency • Loose stools Treatment-Related Risks • Inadequate bilirubin reduction if treatment parameters are not optimized • Prolonged treatment duration requiring escalation of care
2026-02-26 | HTA26008 | Propel | Pioneer Medical Supplies | Intersect ENT | USA | Medical Product/ Kit | Therapeutic | Approved for adoption and for use in patients ≥ 18 years of age following ethmoid sinus surgery to maintain patency, thereby reducing the need for post-operative intervention such as surgical adhesion lysis and/or use of oral steroids. The PROPEL sinus implant separates mucosal tissues, provides stabilization of the middle turbinate, prevents obstruction by adhesions, and reduces edema Recommendations 1. PROPEL® Sinus Implant is recommended for hospital use as an adjunct following endoscopic ethmoid sinus surgery in adult patients (18 years and older) where maintaining sinus patency and reducing post-operative interventions is clinically indicated. 2. The device should be inserted only by trained and competent ENT surgeons under endoscopic visualization, in accordance with the manufacturer’s Instructions for Use. 3. Appropriate patient selection must be ensured, including confirmation of indication and exclusion of contraindications such as known or suspected hypersensitivity to mometasone furoate or to bioabsorbable polymer components. 4. The implant and delivery system are intended for single use only and must not be reused, reprocessed, or re-sterilized; sterility and packaging integrity must be verified prior to use. 5. Proper deployment and handling techniques must be followed, including correct loading, positioning, and atraumatic placement of the implant, without modification or damage to the device. 6. The PROPEL sinus implant should be incorporated into institutional endoscopic sinus surgery postoperative protocols, including routine saline irrigation, scheduled follow-up, and endoscopic assessment as clinically indicated. 7. Initial and periodic staff training and competency validation should be implemented for surgeons and operating room staff involved in implant preparation, placement, and postoperative management. 8. Clinical outcomes and safety events should be systematically monitored and documented as part of routine post-market surveillance and quality reporting, including the need for post-operative interventions, endoscopic findings, unplanned visits or procedures, and any device- or steroid-related adverse events. Clinical Use Overview Population / Intended User: CRS may be broadly defined as an inflammatory disorder of the paranasal sinuses and linings of the nasal passages that lasts12 weeks or longer/ PROPEL sinus implant system is used by physicians in chronic rhinosinusitis (CRS) patients undergoing endoscopic sinus surgery in either the hospital, clinic, or professional office settings. To be Performed By: ENT surgeon / physician (endoscopic placement) Clinical Setting: Hospital Condition of Use: The PROPEL® sinus implant is intended for use in patients ≥ 18 years of age following ethmoid sinus surgery to maintain patency, thereby reducing the need for post-operative intervention such as surgical adhesion lysis and/or use of oral steroids. The PROPEL sinus implant separates mucosal tissues, provides stabilization of the middle turbinate, prevents obstruction by adhesions, and reduces edema Contraindications: 1. Pediatric Use: The safety and effectiveness of the implant in pediatric patients have not been established. 2. Pregnancy and Nursing females. 3. Patients with suspected or confirmed hypersensitivity and/ or intolerance to mometasone furoate. 4. Patients with a known hypersensitivity and/ or intolerance to lactide, glycolide or caprolactone copolymers.
2026-02-26 | HTA26006 | STANDARD™ E TB-Feron ELISA | Babirus Medical Equipment Trading LLC | SD Biosensor, Inc. | South Korea | Diagnostic Test | Diagnostics | Approved for an in vitro diagnostic test using TB-specific recombinant protein Antigens (ESAT-6, CFP-10 and TB 7.7) to stimulate cells in heparinized whole blood. Detection of interferon-gamma (IFN-γ) by enzyme-linked immunosorbent assay (ELISA) is used to identify in vitro responses to those Recombinant TB Antigens that are associated with Mycobacterium tuberculosis infection. The test is an sandwich test for M. tuberculosis infection (including disease) and is intended for use in conjunction with risk assessment, radiography, and other medical and diagnostic evaluations. • Use TB-Feron ELISA in laboratories that already have basic ELISA capability, including incubation, centrifugation, plate washing, and result reporting systems. • Clinical decision-making should follow existing TB guidelines, including further evaluation of positive results and cautious interpretation of negative results in high-risk patients. • A clear plan should be in place to manage borderline or indeterminate results, such as repeat testing or clinical follow-up. • Monitor indeterminate result rates, especially in high-risk groups such as immunocompromised patients, people living with HIV, and children. • Implement routine quality control checks using kit controls and internal tube controls. • Ensure proper chemical safety measures are in place when handling reagents such as acidic stop solution, including PPE use, spill procedures, and appropriate waste disposal. Clinical Overview: Population: Individuals undergoing TB infection screening, especially high-risk groups To be performed by: Trained medical laboratory technologist or technician Clinical setting: Hospital, clinic, or reference laboratory that are approved for TB diagnosis Condition of use: Used as an IGRA test for TB infection following manufacturer instructions, with further clinical evaluation if results are positive.
2026-02-26 | HTA26005 | PANA RealTyper™ HPV Kit | Babirus Medical Equipment Trading LLC | HLB Panagene Co., Ltd. | South Korea | Diagnostic Test | Diagnostics | Approved for in vitro diagnostic reagent for genotyping of human papilloma virus (HPV) using peptide nucleic acid (PNA) probes. This kit is an amplified DNA test for the qualitative detection of a total of 40 HPV genotypes in a real-time PCR (polymerase chain reaction) system. This kit provides genotyping information of 20 high-risk and 2 low-risk types using melting temperature (Tm) analysis. Furthermore, this kit also detects 18 other genotypes (without genotyping) in DNA samples from specimens. • The PANA RealTyper™ HPV Kit may be considered as an adjunct molecular genotyping assay, particularly where extended HPV genotype information is required. • Use should be supported by local laboratory verification and alignment with existing HPV testing workflows. • Results should be interpreted with clear reporting to avoid unnecessary clinical intervention Clinical Overview Population: Females undergoing HPV testing Performed by: The PANA RealTyper™ HPV Kit is to be used by trained laboratory professionals or molecular lab technologist within a clinical laboratory environment only. Clinical Setting: Accredited clinical molecular diagnostics laboratory Condition of Use: PCR-based HPV DNA testing with validated extraction, quality controls, and guideline-aligned result interpretation.
2026-02-26 | HTA26003 | Electronic Fetal Monitor \ Fetal Lite NST-CTG | Emirates & World Medical Supplies | InnAccel Technologies Pvt. Ltd. | India | Medical Device | Diagnostics | Approved for adoption and intended to be used non-invasively to measure and monitor Fetal Heart Rate (FHR) & Uterine Contractions (UC) during pregnancy (>24 weeks) and during labour. Indications of Use The Fetal Lite Electronic Fetal Monitor is intended to be used on pregnant women post 24 weeks of gestation. Recommendations 1. Fetal Lite NST-CTG is recommended for hospital use for antenatal and intrapartum fetal monitoring where cardiotocography is clinically indicated. 2. The device should be operated by trained and competent healthcare professionals only. 3. The device is an adjunct tool and must not replace clinical judgment; consider all clinical information. 4. Continuous maternal and fetal clinical assessment must accompany CTG monitoring and automated interpretation outputs. 5. Proper probe placement, signal verification, and routine device performance checks must be ensured. 6. The system should be incorporated into obstetric unit NST and labor ward fetal monitoring protocols. 7. Initial and periodic staff training and competency assessment should be implemented. 8. Clinical outcomes and safety events should be monitored and documented as part of routine quality assurance and post-market surveillance. Clinical Use Overview Population / Intended User Pregnant women ≥24 weeks gestation\ monitoring of fetal heart rate and uterine contractions in pregnant women from 24 weeks gestation, during antenatal surveillance and intrapartum care. To be Performed By Obstetricians, gynecologists, trained nurses, and midwives Clinical Setting Hospitals, maternity units, antenatal clinics, labor wards Condition of Use Use in accordance with manufacturer IFU under trained clinical supervision Contraindications • The Fetal Lite Electronic Fetal Monitor should not be used on patients with abdominal skin rashes at the intended probe placement site. • The device should not be used on patients with known skin allergies at the intended probe placement site. • The device should not be used on patients with scarring at the intended probe placement site that may interfere with signal acquisition. • The device should not be used on patients with dermatitis at the intended probe placement site. • The device should not be used on patients with open wounds or cuts on the abdomen where probes are to be applied.
2026-02-26 | HTA26002 | VAPCare – Automated Secretion Clearance and Oral Hygiene Management System Model F1-01-000-00-000 | Emirates & World Medical Supplies | InnAccel Technologies Pvt. Ltd. | India | Medical Device | Pulmonology Critical Care Infectious Disease Prevention | Approved for Adoption and it’s intended for use in adults, to remove secretions and other fluids from above the endotracheal tube cuff and provide oral lavage or irrigation. Recommendations 1. VAPCare is recommended for hospital use in adult ICU patients requiring invasive mechanical ventilation for more than 48 hours where VAP prevention is clinically indicated. 2. VAPCare can be used as adjunct to mechanical ventilation care, enhancing existing VAP prevention bundles by automating secretion clearance and oral hygiene. 3. To be operated only by trained and competent healthcare professionals 4. Strict adherence to infection prevention and disposable handling protocols must be ensured 5. Continuous patient monitoring, including respiratory status and mucosal integrity, must be maintained 6. Device settings, including suction pressure and frequency, should be optimized according to institutional protocols and patient condition 7. VAPCare should be incorporated into ICU ventilator-associated pneumonia prevention bundles 8. Initial and periodic staff training and competency assessment should be implemented 9. Clinical outcomes, including VAP incidence, duration of ventilation, antibiotic usage, and safety events, should be monitored and documented as part of post-market surveillance and quality reporting Clinical Use Overview Population / Intended User Adult mechanically ventilated ICU patients\ VAPCare is indicated for adult patients requiring invasive mechanical ventilation for more than 48 hours, to provide sensor-based secretion clearance and oral hygiene management in the oral, oropharyngeal, and subglottic regions, including removal of secretions above the endotracheal tube cuff. To be Performed By Trained ICU nurses, respiratory therapists, and physicians Clinical Setting Intensive Care Units (ICU) Condition of Use Invasive mechanical ventilation >48 hours Contraindications 1. Conscious patients with an intact gag reflex 2. Severe orofacial trauma or unstable facial fractures 3. Severe laryngospasm or stridor 4. Acute severe epiglottitis or croup 5. Severe cardiovascular instability where suctioning may exacerbate hemodynamic compromise 6. Patients unable to tolerate oral or oropharyngeal device placement 7. Patients with conditions where oral or subglottic suctioning is clinically contraindicated per treating physician judgment
2026-02-26 | HTA26001 | VOLT PFA System | Abu Dhabi International Medical Services | Abbott Medical | USA | Medical Device | Therapeutic | Approved for the indication for the treatment of symptomatic, recurrent, drug-refractory paroxysmal or persistent (episode duration less than one year) atrial fibrillation when used in conjunction with a compatible pulsed field ablation (PFA) generator. The catheter is compatible with the EnSite™ X EP System. The Current™ PFA Generator is indicated for use with compatible ablation catheters for the treatment of cardiac arrhythmias. The Volt™ PFA Catheter, Sensor Enabled™ is intended for use in patients for cardiac electrophysiological mapping, pacing, and ablation. • Use should be limited to fully equipped electrophysiology laboratories and performed by trained, experienced electrophysiologists. • Strict compliance with the Instructions for Use and ongoing local outcome monitoring are recommended to ensure safety and quality. Clinical Overview: Population: The Volt™ PFA Catheter, Sensor Enabled™ is intended for patients with atrial fibrillation undergoing cardiac ablation procedures who have a vascular or cardiac anatomy compatible with an introducer with a minimum inner diameter of 12.5 F. Performed by: The Volt™ PFA Catheter, Sensor Enabled™ is intended for use only by physicians thoroughly trained in the techniques of catheter ablation. Clinical Setting: a fully equipped electrophysiology laboratory. Condition of Use: Pulmonary vein ablation using the Volt™ PFA balloon catheter, performed through a vein access to the heart, guided by a mapping system, under anesthesia or sedation, with blood-thinning medication, following the manufacturer’s instructions.
2026-02-26 | HTA25087 | Picterus Jaundice Pro | Madia Trading and Investment FZCO | Picterus AS | Norway | System/ Software/ AI / App | Screening | Conditionally Approved with RWE generation- The Picterus® Jaundice Pro is intended to assist in screening and follow up of neonatal jaundice. Moreover, mandatory confirmatory testing before treatment decisions is a must. The medical device is intended to support authorized healthcare providers, parents and/or caregivers (lay persons) of newborns in the assessment of neonatal jaundice. The medical device is intended for use as a screening and control device together with other clinical signs and laboratory measurements. Assessments shall never be made using the Picterus® Jaundice Pro and the Picterus® Calibration Card alone. In cases where the bilirubin values from the Picterus® Jaundice Pro app indicate hyperbilirubinemia, the specific bilirubin levels of the newborn must always be confirmed by other methods, such as serum bilirubin, before treatment is established. The Picterus® Jaundice Pro and the Picterus® Calibration Card are designed for use in clinical and home settings. It should NOT be used for: • Replace laboratory bilirubin for treatment decisions • Be used as sole diagnostic tool in high-risk neonates • mandatory confirmatory testing before treatment decisions. Recommendation: • Adoption in Postnatal Care: To introduce this device in both hospital post-discharge follow-up and community settings ( home visits, health visitor check-ups). Its use should complement existing protocols – e.g. as a screening step when infants are visibly jaundiced or at high risk, to decide if a hospital referral for TSB is needed. • Training and User Support: Ensure comprehensive training for all users. Neonatal nurses and community healthcare workers should be trained to use the app and card correctly, interpret the results, and educate parents. Parents who participate should receive supervised instruction (e.g. before discharge, a midwife can demonstrate the app on the newborn) . Clear educational materials and in-app tutorials (provided by the manufacturer) should be utilized to maximize correct usage and confidence. Ongoing support (helpdesk or digital resources) will further facilitate proper use by caregivers at home. • Clinical Protocol Integration: Update guidelines to incorporate smartphone bilirubin screening. NHS trusts should develop or update local neonatal jaundice protocols to include Picterus JP as an option for screening. Define thresholds and actions. • Quality Assurance & Maintenance: Implement QA procedures and device maintenance protocols. Assign responsibility (e.g. a neonatal nurse lead or clinical physics team) for regular quality checks of the Picterus system. • Data Management and Oversight: Establish a clear process for result documentation and follow-up. When Picterus JP is used by parents at home, there should be a mechanism to report the results to a healthcare provider (e.g. through a connected app portal or by phone. • Monitoring and Feedback: Gather user feedback and outcomes to inform future use. As the technology is adopted, collect data on outcomes – how many referrals were made, how many true positives vs false positives, any cases of missed high bilirubin, etc. Also gather feedback from both healthcare staff and parents on usability and satisfaction. This information will be valuable for refining training and possibly working with the manufacturer on any software improvements. Additional recommendations: • Mandatory confirmatory serum testing before phototherapy • Ongoing audit of: o False negative rate o Referral rate o Hospital readmission rate Population/Intended User: • Newborn infants Term and near-term neonates ≥35 weeks gestation, Clinically stable infants, Early discharge populations, primary care follow-up infants Use cautiously in: • Preterm <35 weeks, Hemolytic disease, G6PD deficiency, Rapidly rising bilirubin, Ill neonates. To be performed by: healthcare providers, parents and/or caregivers Settings: • Maternity wards • Postnatal discharge clinics • Rural / remote primary care • Telehealth follow-up programs • Home settings Not recommended as primary tool in: • NICU for critically ill neonates • Exchange transfusion decision-making Condition of use: Use in accordance with manufacturer IFU under trained clinical supervision Contraindication: Do NOT use the app on newborns with: a. Jaundice appearing within the first 24 hours of life. b. Current or previous treatment involving phototherapy (light therapy) or prolonged exposure to sunlight because safety assessment for this patient group has not been performed. Light can alter the relationship between bilirubin levels in the blood and skin, affecting the accuracy of the result. c. A condition or treatment that affects skin color or appearance, excluding jaundice, such as skin diseases or extremely dry skin that appears cracked or peeled. The skin must be intact. Additionally, do NOT use the device for monitoring newborns at high risk for rapidly rising bilirubin levels due to confirmed or suspected underlying conditions, including but not limited to: - Hemolytic disease. - G6PD deficiency. - Cephalohematoma or birth trauma. - Red blood cell disorders. Conditionally approved for RWE follow up Centres adopting Picterus should collect: 1. Percentage of newborns screened 2. Reduction in serum bilirubin blood draws 3. Readmission rate for hyperbilirubinemia 4. Cases of severe hyperbilirubinemia missed 5. Parent satisfaction scores Reassessment after 1-2 years. Conditionally approved as per AAMEN Team: Right to audit by regulatory body
2026-01-26 | HTA25118 | Plasma Surgery System PLA-700 | RIGHTVIEW MEDICAL EQUIPMENT TRADING LLC | Chengdu Mechan Electronic Technology Co. Ltd. | CHINA | Medical Device | Therapeutic | Approval for Adoption for ablation and coagulation of soft tissues, including the throat, nose, joints, ligaments, prostate, and other areas. Its designed for otolaryngological (ENT) surgery. It is suitable for procedures involving the oral cavity, throat, and endoscopic sinus surgeries, and can be used for both pediatric and adult patients, regardless of gender. Clinical Use Overview: Population / Intended User: PLA-700 is targeted to kid or adult, male or female and intended for ablation and coagulation of soft tissues, including throat, nose, joint, ligiments and prostate To be Performed By: The operator of Plasma Surgery System must be a physician or medical personnel under the supervision of a physician and must have received sufficient training in clinical electrosurgical techniques. Clinical Setting: The device is intended to be used in operating room as well as qualified hospital environment. Contraindications: The use of this device is contraindicated in patients with the following conditions: a) Patients with cardiac, renal and brain insufficiency, severe anemia, or bleeding tendency. b) Patients who are using physiological monitor with needle style instrument. c) Patients with combined acute infection. d) Patients with cruor dysfunction, severe infection, or malignant diseases. e) Patients who are using heart pacemaker or pacemaker electrode can’t be treated by this device and also can’t go near when it works. f) Patients embedded with metallic implants. g) Patients with acute infection can’t be treated by this device unless be treated with anti-inflammatory and infection control at first. h) Patients during gestation. Recommendations: PLA-700 is a bipolar, radiofrequency electrosurgical system. PLA-700 is provided for otolaryngological (ENT) surgery, including the oral, throat and endoscopic sinus surgical procedures, targeted to kid or adult, male or female. Based on the available clinical evidence, manufacturer-provided clinical trial data, and publicly available literature, the PLA-700 Plasma Surgery System may be considered appropriate for adoption in ENT surgical procedures requiring soft tissue ablation, cutting, coagulation, and hemostasis, subject to the following conditions: 1. Approval for Adoption for ablation and coagulation of soft tissues, including the throat, nose, joints, ligaments, prostate, and other areas. Its designed for otolaryngological (ENT) surgery. It is suitable for procedures involving the oral cavity, throat, and endoscopic sinus surgeries, and can be used for both pediatric and adult patients, regardless of gender. 2. Use should be restricted to qualified healthcare facilities equipped with appropriate operating room infrastructure and in compliance with local regulatory and electrical safety requirements. 3. The device should be operated only by trained physicians or medical personnel under physician supervision, who have received formal training in electrosurgical and plasma-based techniques, in accordance with the manufacturer's instructions for use (IFU). 4. Adoption should be limited to approved clinical indications, including otolaryngological (ENT) procedures involving the oral cavity, throat, and endoscopic sinus surgery, as specified in the IFU. 5. Patient selection criteria and contraindications outlined by the manufacturer must be strictly adhered to, including precautions for patients with implanted electronic devices (e.g., pacemakers), severe systemic illnesses, active infections, bleeding disorders, or pregnancy. 6. The device should be used exclusively with compatible single-use plasma probes and accessories, and the availability of required consumables must be ensured prior to clinical use. 7. Healthcare institutions should establish and follow local protocols for perioperative safety, infection control, and adverse event reporting. 8. Ongoing post-market surveillance and incident reporting should be conducted in accordance with institutional policies and regulatory requirements.
2026-01-26 | HTA25116 | Plasma Surgery System PLA-800 | RIGHTVIEW MEDICAL EQUIPMENT TRADING LLC | Chengdu Mechan Electronic Technology Co. Ltd. | CHINA | Medical Device | Therapeutic | Approved The Plasma Surgery System PLA-800 is intended as an orthopedics-purposed electrosurgical generator designed for ablation with normal saline solution and coagulation to be delivered to soft tissue through a disposable plasma probe during endoscope surgical procedures. This device is intended for use by qualified personnel and is for use only with appropriate Mechan accessories. Osteoarthritis, meniscus injury, ligamentous injury, cartilage injury, adhesive capsulitis, and degenerative spine disease. • Ensure availability of required single-use probes. • Verify standards compliance, enforce IFU training and maintain post-market incident reporting. Population: for orthopedic surgery To be performed by: The operator of Plasma Surgery System must be a physician or medical personnel under the supervision of a physician and must have received sufficient training in clinical electrosurgical techniques. Clinical setting: The device is intended to be used in operating room as well as qualified hospital environment. Condition of use: The use of this device is contraindicated in patients with the following conditions: a) Patients with cardiac, renal and brain insufficiency, severe anemia, or bleeding tendency. b) Patients who are using physiological monitor with needle style instruments. c) Patients with combined acute infection. d) Patients with cruor dysfunction, severe infection, or malignant diseases. e) Patients who are using heart pacemaker or pacemaker electrode can’t be treated by this device and also can’t go near when it works. f) Patients embedded with metallic implants. g) Patients with acute infection can’t be treated by this device unless they be treated with anti-inflammatory and infection control at first. h) Patients during gestation. i) For special surgery please refer to the instructions to learn more details of contraindications. Patients with Pacemakers Use with caution in the presence of internal or external pacemakers. Interference from an electrosurgical current can cause a pacemaker to enter an asynchronous mode or can block the pacemaker effect entirely. For further information, consult the pacemaker manufacturer or hospital Cardiology Department.
2026-01-26 | HTA25115 | PATHFAST POCT Analyzer - PATHFAST Presepsin | Al Shomoukh Trading for Technical & Medical Supplies – ASTEMED L.L.C | Pathfast | Japan | In vitro diagnostic (IVD) analyzers and associated test kits | Diagnostics | Approved for - as an aid in the early detection and differential diagnosis of clinically relevant bacterial infections, - as an aid in the assessment of the degree of septic severity and in the risk stratification of patients with systemic bacterial infection, sepsis, severe sepsis and septic shock, - as an aid in decision making on antibiotic therapy for patients with lower respiratory tract infections and patients with suspected or confirmed sepsis, - by laboratory technician, nurse or physician, at point of care (POC). Clinical Use Overview: Population / Intended User: Patients with suspected sepsis or systemic inflammatory response syndrome (SIRS) To be Performed By: Doctors, Nurses, Trained, competency-validated operators Clinical Setting: Hospital, ED, ICU, POC Condition of Use: PATHFAST Presepsin is a product for in-vitro diagnostic use with the in vitro diagnostic (IVD) automated analyser PATHFAST for the quantitative measurement of presepsin in human heparinized or EDTA whole blood and plasma. PATHFAST Presepsin is intended to be used: - as an aid in the diagnosis and prognosis of sepsis, - in the assessment of the degree of septic severity, - in the risk stratification of critically ill septic patients, - by laboratory technician, nurse or physician, - in hospital including emergency room, doctor’s office and clinical laboratory. P ATHFAST Presepsin is a device for near patient testing (NPT). Contraindications: • No patient contraindications for the analyzer itself; limitations are assay-specific (e.g., incorrect specimen type, improper handling, out-of-spec samples). • Follow each assay IFU/package insert. Recommendations: 1. Adopt PATHFAST Presepsin testing for early assessment, severity stratification, and prognostic support in patients with suspected sepsis or SIRS, specifically within ED, ICU, and acute care hospital settings where rapid clinical decision-making is required. 2. Integrate Presepsin testing as an adjunct biomarker within established sepsis management pathways, alongside clinical assessment, vital signs, lactate levels, cultures, and other relevant laboratory markers. 3. Restrict operation of the device to trained and competency-validated healthcare professionals, such as laboratory staff, nurses, or physicians, in accordance with manufacturer IFU and institutional policies. 4. Perform local verification of the Presepsin assay against the site’s reference laboratory method prior ti define analytical performance limits and clinical cut-offs consistent with IFU and clinical protocols. 5. Ensure routine internal quality control and participation in external quality assessment with mandatory lot-to-lot verification and calibration upon reagent changes, and a documented process for recall, FSN tracking, and clinician notification. 6. Use Presepsin results strictly for the intended clinical indications. 7. Apply manufacturer-recommended cleaning and disinfection procedures and comply with institutional infection prevention and biomedical waste management policies, especially in high-risk acute care environments.
2026-01-26 | HTA25114 | Implantable Pectus-up Surgery Kit - model 02-0002 | Al Zahrawi Medical Supplies, LLC | Ventura Medical Technologies, SL | Spain | Product/ Kit | Therapeutic | Conditional approval of the Pectus-Up Implant System for use solely within hospitals or specialized surgical centers performing chest wall deformity correction in patients with pectus excavatum who meet specified clinical and anatomical criteria, including cases of pectus excavatum and other similar rib cage deformities. In addition, to report to DoH the REW data on an annual basis regarding the effectiveness and efficacy of the technology. Clinical Use Overview: Population / Intended User: Candidates are individuals with pectus excavatum who fall within an optimal age range of ~12–30 years, where the chest wall is sufficiently pliable for safe, effective sternal elevation To be Performed By: Surgeon Clinical Setting: Hospital Condition of Use: The Implantable Pectus Up Surgery Kit is indicated for the remedial/corrective treatment of patients with Pectus Excavatum. Candidates are individuals with Pectus Excavatum who meet the following criteria: • Age Range: Typically between approximately 12 and 30 years old, as the chest wall in this age group is generally sufficiently pliable for safe and effective sternal elevation. Patients over 30 years can still be treated; however, they often require adjunct techniques (such as chondrectomy or a modified Ravitch procedure) to address increased rigidity. • Anatomical Severity: Patients should exhibit severity levels consistent with intervention based on imaging. While the Haller Index helps stratify severity—>3.2 indicating severe deformity, 2.5–3.2 for mild to moderate, and <2.5 usually not requiring intervention—it is not an absolute criterion. Clinical judgment, considering other indices and symptoms, remains paramount in determining candidacy. • Sternal Anatomy: Adequate sternal anatomy is essential for implant anchoring, including a minimum sternal diameter of ≥28 mm at the intended plate site. Markedly irregular sternal surfaces may necessitate additional supports or combined techniques. • Bone Quality: Patients should have sufficient sternal bone quality, typically reflected by CT Hounsfield Units (HU) indicative of secure fixation—generally >100–300 HU for cortical bone, with >200 HU expected in healthy, younger patients. Those with osteopenia, osteoporosis, or on bone-weakening therapies (e.g., steroids) may be unsuitable for Pectus Up alone and should be considered for combined approaches. • Chest Wall Rigidity: Patients should not exhibit prohibit rigidity of the chest wall due to cartilage calcification or ossification. If rigidity is present—more common in older patients—adjunct procedures such as chondrectomy can restore flexibility and enable safe elevation. An individualized assessment including CT-based evaluation of bone density and morphology guides patient selection and determines whether the device can be used alone or in combination with other techniques to optimize stability and outcomes. Contraindications: Patients Contraindications • Inability or unwillingness to follow instructions (due to mental or neurologic illness) • Insufficient bone quality or quantity for fixation or remodeling • Active infection at the surgical site • Metal allergy or sensitization (exclusion criterion) Warnings and Precautions • The device provides structural support but does not replace native chest-wall structures. • The amount of corrections achievable cannot be precisely predicted. • Meticulous preoperative planning is essential, including correct sizing and positioning. • Avoid placement of screws in the xiphoid process. • Ensure appropriate screw length to prevent complications. • The device is intended for single use only; reuse increases the risk of failure and infection. • MRI procedures are not recommended due to unknown safety. • Postoperative activity restrictions include avoiding heavy loads and contact sports early after surgery to prevent loosening or breakage. • Possible adverse effects include: o Metal sensitivity or allergy o Pain or discomfort o Skin irritation, seroma, or infection o Implant fracture, loosening, or displacement o Incomplete correction or relapse o Risks associated with general anesthesia Healthcare Practitioners Risks and Controls • The system is supplied non-sterile and requires validated hospital steam sterilization, following specific cycle parameters to prevent contamination and ensure sterilization effectiveness. • Standard operating room hazards (e.g., sharps, instrumentation) are addressed through proper training, adherence to the Instructions for Use (IFU), and risk-control measures. • All risk management procedures and measures are documented in the risk management file. Recommendations: 1. Conditional approval of the Pectus-Up Implant System for use solely within hospitals or specialized surgical centers performing chest wall deformity correction in patients with pectus excavatum who meet specified clinical and anatomical criteria, including cases of pectus excavatum and other similar rib cage deformities. 2. Use should be restricted to trained surgeons with demonstrated experience in chest wall deformity correction and documented, device-specific training on the Pectus-Up implant system. 3. Strict patient selection criteria must be applied, including: o Confirmed diagnosis of pectus excavatum o Symptomatic or functionally significant deformity o Adequate chest wall anatomy and sufficient sternal bone quality o Absence of contraindications (e.g., active infection, metal allergy) 4. Preoperative assessment must include chest CT imaging with calculation of thoracic indices (e.g., Haller Index) and evaluation of chest wall rigidity, bone density, and morphology to confirm suitability for implantation. 5. The procedure must be performed in accordance with the manufacturer’s Instructions for Use (IFU), including appropriate implant sizing, positioning, fixation technique, and perioperative risk mitigation. 6. Perioperative and postoperative monitoring protocols should be implemented to assess surgical complications, pain control, functional recovery, and adequacy of correction, with documented follow-up at defined clinical intervals. 7. The device should be considered a temporary implant, with planned removal once adequate chest wall remodeling is achieved, typically within 2–3 years, or as clinically determined by the treating surgeon. 8. Post-market clinical follow-up (PMCF) and local data collection are recommended to capture: o Degree and durability of correction o Implant-related complications or reintervention o Safety profile in the local healthcare setting 9. Clinical outcomes and safety events should be monitored, documented with reports submitted to the DoH PV or real-world data team for long-term use assessments and in case of concerns or safety issues. 1 0. To report to DoH the REW data on an annual basis regarding the effectiveness and efficacy of the technology.
2026-01-26 | HTA25113 | STEPHANIE VENTILATOR | HESSA MEDICAL EQUIPMENTS CO LLC | Fritz Stephan GmbH | Germany | Medical Device | Therapeutic | Approved and intended for use in providing short-term or long-term ventilation and ventilatory support in cases of respiratory insufficiency. Intended medically for invasive and non-invasive ventilation in an ICU setting, as well as for high flow O2 therapy, and is suitable for long-term ventilation. STEPHANIE is used as an ICU ventilator and can be used for transport within the hospital. The STEPHANIE is suitable for the ventilation of preterm infants and newborns, children, and adults, and may only be used by authorized specialist personnel. Clinical Use Overview: Population / Intended User: neonates, pediatric, adults To be Performed By: The ventilator may only be used by authorized personnel such as physicians, qualified medical specialists, nurses and respiratory therapists. Clinical Setting: Hospital Condition of Use: The STEPHANIE is intended for use in providing short-term or long-term ventilation and ventilatory support in cases of respiratory insufficiency. Contraindications: Non-invasive ventilation is contraindicated in the following cases: ▪ No spontaneous breathing ▪ Fixed or functional airway obstruction ▪ Gastrointestinal bleeding or ileus ▪ Uncontrolled vomiting ▪ Inability to protect the airway ▪ Skull fractures ▪ Patient deterioration Recommendations: The STEPHANIE ICU Ventilator is recommended for hospital use in NICU, PICU, and ICU settings for the provision of invasive and non-invasive mechanical ventilation where clinically indicated. 1. The ventilator should be operated only by trained and competent healthcare professionals, including physicians, nurses, and respiratory therapists. 2. Appropriate patient monitoring—including oxygenation, ventilation parameters, hemodynamic status, and alarms—must be maintained throughout ventilator use. 3. Ventilation settings should be selected and adjusted according to the patient’s clinical condition and in accordance with institutional respiratory care protocols. 4. Lung-protective ventilation strategies should be applied, including: o Proper control of airway pressures and tidal volumes. o Use of advanced ventilation modes (e.g., volume-targeted ventilation or HFOV) where clinically indicated. 5. The device should be used in accordance with the manufacturer’s instructions for use (IFU), including environmental, gas supply, and accessory requirements. 6. The STEPHANIE ventilator should be incorporated into ICU/NICU respiratory care protocols, including escalation and de-escalation pathways. 7. Initial and periodic staff training, competency assessments, and device familiarization should be ensured. 8. Preventive maintenance, calibration, and safety checks must be performed by biomedical engineers in accordance with the manufacturer’s recommendations. 9. Clinical outcomes and safety events (e.g., duration of ventilation, complications, need for escalation, and adverse events) should be monitored and documented as part of routine post-market surveillance and quality assurance activities.
2025-12-26 | HTA25111 | Rotem Sigma | Nafees and Masoud Medical supplies | Werfen | Spain | Medical Device | Diagnostics | Approved -ROTEM sigma thromboelastometry system is a fully integrated and automated in vitro diagnostic system designed to monitor and analyze a patient's coagulation status by measuring the viscoelastic properties of a 3.2% citrated venous or arterial whole blood sample. • Training- Ensure structured training and competency assessment for all operators and clinicians interpreting results. • Quality-Implement routine internal quality control, maintenance, and participation in external quality assurance where available. • Population: adult patients 21 years and older • To be performed by: Trained healthcare professionals • Clinical setting: healthcare facilities • Condition of use: Point-of-care testing using citrated whole blood, alongside clinical assessment and other laboratory tests
2025-12-26 | HTA25112 | SURFACER INSIDE-OUT | ARABIAN MILLENNIUM MEDICAL TRADING LLC | MERIT MEDICAL USA | USA | Medical Device | Therapeutic | Approved to obtain central venous access to facilitate catheter insertion into the central venous system for patients with upper body venous occlusions or other conditions that preclude central venous access by conventional methods. Recommendations: 1. Approve the Surfacer® Inside-Out® Access Catheter System for use in tertiary care hospitals as a specialized solution for patients with complex upper-body central venous occlusions where conventional access methods have failed. 2. Restrict use to trained interventional specialists in facilities equipped with fluoroscopic imaging, emergency surgical support, and blood transfusion services. 3. Mandate strict adherence to the manufacturer’s Instructions for Use (IFU), including single-use only (no reuse or re-sterilization), patient selection criteria, contraindications, and post-procedure monitoring, with institutional oversight for safety and quality assurance. 4. Institutions should monitor safety outcomes and procedural success as part of quality assurance. Population: Adult patients (≥18 years) requiring central venous access with upper-body venous occlusions To be Performed By: Physicians/ Surgeons experienced in interventional procedures Clinical Setting: Hospitals with fluoroscopic capability where emergency surgery and blood transfusion services are available. Condition of Use: The Surfacer Inside-Out Access Catheter System is intended to obtain central venous access to facilitate catheter insertion into the central venous system for patients with upper body venous occlusions or other conditions that preclude central venous access by conventional methods. Contraindications: 1. The Surfacer® System is contraindicated for patients with an occlusion of the right femoral vein, right iliac vein or inferior vena cava, or acute thrombosis within any vessel to be crossed by the Surfacer® System. Special precautions may be required for patients with coagulation disorders or on anti-coagulation therapy. 2. The Surfacer® System is not intended for use in the coronary or cerebral vasculature. 3. This device is not to be used in the arterial system.
2025-12-26 | HTA25110 | EyeLib Device | S A W A MEDTEC MEDICAL EQUIPMENT TRADING LLC | MIKAJAKI SA | France | Medical Device | Diagnostics | Conditionally Approved : The aim of the EyeLib station is to combine several diagnostic medical devices into one platform to streamline the testing process by improving time efficiency, reducing demand of operator time, reducing required level of operator training, and enhancing the overall workflow of the healthcare facility. Consider adoption of EyeLib™ Suite for ophthalmology screening and triage in Abu Dhabi government healthcare settings to enhance efficiency and patient access. Monitor implementation outcomes and support further research to validate comparative effectiveness and cost savings. Ensure ongoing compliance with regulatory and quality standards. This system is intended to support early detection, triage and monitoring of common ocular diseases while reducing dependencies on specialist aviablity for first line assessment. It will not replace the final clinical decisions which remains with eyecare professionals. Conditionally Approved as per AAMEN: remote access to the device to be disabled. Population: For vision screening, ocular monitoring especially for patient who require frequent eye tests with diabetes, high risk groups and elderly, To be performed by: Ophthalmologists, Optometrists, Opticians, Orthoptists or other trained health personnel. Clinical Setting: Clinics, Hospitals and ophthalmological Centers. Condition of use: EyeLib operations training given by a MIKAJAKI qualified person is mandatory before the use of the station. Contraindication: Do not use the EyeLib for those patients who: • Have any trouble history related to photodermatitis. • Have undergone photodynamic therapy (PDT).
2025-12-26 | HTA25107 | PillSense System | Mart Medical Solution | EnteraSense Ltd | Ireland | Medical Device | Diagnostics | Condtionally approved as: The PillSense™ System is a prescription only device consisting of a reusable receiver and single-use ingestible capsule, intended to be used for the detection of blood in the upper gastrointestinal tract in hemodynamically stable adults suspected of having upper gastrointestinal bleeding (UGIB). Consider as an adjunct triage tool for suspected upper GI bleeding. Not replacmement for endoscopy; results must be clinically interpreted. Requires clear clinical protocols and trained staff for safe implemenntation. • Access to imaging: Although capsule retention was rare, facilities need access to abdominal X‑ray or other imaging to confirm passage or locate the capsule if symptoms arise. Endoscopic retrieval may be required for retained capsules. • Electronic interference: The device must be used away from strong electromagnetic fields and patients with implanted electromedical devices are excluded. Facilities should review compatibility with other equipment. Population: stable adults suspected of having upper gastrointestinal bleeding To be performed by: Healthcare professionals, gastroenterologists. Clinical Setting: Healthcare facilities Condition of use: consider Pill Sense as an adjunct for UGIB triage while advocating for further research and ensuring facilities have the necessary protocols and infrastructure for safe implementation. Condtionally approved as per AAMEN: Replaced Receiver to be managed as per the facility asset disposal policy.
2025-12-26 | HTA25100 | maviLED Phototherapy System | HESSA MEDICAL EQUIPMENTS CO LLC | ASTEK MEDICAL | TURKEY | Medical device | Therapeutic | Approved for Treatment of neonatal hyperbilirubinemia (unconjugated jaundice). Recommendations: 1. maviLED Phototherapy System is recommended for hospital use in the treatment of neonatal hyperbilirubinemia where phototherapy is clinically indicated 2. To be operated by trained and competent healthcare professionals only 3. Eye protection must be applied throughout phototherapy. 4. Continuous neonatal monitoring (temperature, hydration, and clinical status) must be ensured. 5. Dose delivery must be optimized, including: - Maintaining recommended working distance (30–50 cm). - Adjusting light intensity according to treatment protocol. 6. To be included in NICU and newborn unit phototherapy protocols. 7. Initial and periodic staff training and competency validation should be ensured. 8. Clinical outcomes (TSB reduction rate, duration of therapy, need for escalation) and safety events should be monitored and documented as part of routine post-market surveillance and quality reporting. Clinical Use Overview: Population / Intended User: Neonates To be Performed By: Doctors/ Nurses (use carried out by trained clinical staff under medical supervision) Clinical Setting: Hospitals (NICU, newborn nursery) Condition of Use: maviLED device is a floor-stand, mobile, phototherapy light intended for the treatment of neonatal hyperbilirubinemia. Contraindications: • Congenital porphyria or known photosensitivity disorders. • Neonates receiving photosensitizing medications, unless clinically justified. • Conjugated (direct) hyperbilirubinemia, where phototherapy is ineffective; underlying cause should be investigated. • Use with caution in clinically unstable neonates where phototherapy may interfere with essential monitoring or care, based on physician judgment.
2025-12-26 | HTA25099 | INDIBA ACTIV Therapy System CT8-AT7 | Al Zahrawi Medical Supplies LLC | INDIBA | Spain | Medical Device | Therapeutic | Approved- CT8, AT7 Recommendations • Adoption: Consider INDIBA CT8 for physiotherapy and rehabilitation centers where post-injury pain and functional recovery are common. It may be especially useful in sports medicine (e.g. tendon/ligament injuries), chronic pain clinics (e.g. pelvic pain), and medical settings. Use it adjunctively – combine with manual therapy or exercises for best effect. • Training: Ensure operators have formal training in TECAR therapy techniques. The manufacturer offers specialized courses, and clinicians should learn proper electrode handling, power settings (e.g. use of Indiba Analog Scale), and gel application. • Equipment Checks: Implement regular quality control: calibrate the RF output per manufacturer guidelines, inspect electrodes for wear, and replace gel that is expired. Follow cleaning/disinfection protocols for electrode handles and returns. • Safety Protocols: Screen patients for contraindications (implants, pregnancy, cancer, etc.) before treatment. Monitor vital signs (blood pressure) during initial sessions, especially with spinal or large-joint applications. Educate patients about transient effects (e.g. mild post-treatment soreness) and how to manage them. • Evidence Tracking: Collect routine outcomes data (pain scores, function) to audit effectiveness in your setting. Align treatment protocols (session duration, intensity levels, number of sessions) with those in published studies to maximize chances of benefit. Population: Adults with musculoskeletal injuries (e.g. tendonitis, arthritis, back/neck pain), chronic pain syndromes (e.g. pelvic pain), and patients seeking non-invasive functional rehabilitation. PHYSIOTHERAPY/REHABILITATION IN ADULTS • Injuries of the musculoskeletal system (muscle contracture, tears and strains, tendinitis, sprains, synovitis, bursitis). • Treatment of painful joint conditions (arthritis, osteoarthritis). • Inflammatory processes accompanied by pain in the extremities and spine (neck pain, back pain, lumbago, sciatica,herniated disc, carpal tunnel syndrome, tension headaches). • External treatment of musculoskeletal pelvic floor injuries. • External treatment of chronic pelvic pain Performed by: Healthcare Professionals and trained physiotherapists, Clinical setting: Healthcare facilities, physiotherpy clinics, rehabilitation units
2025-12-26 | HTA25097 | Genedrive MT-RNR1 Test | Genedrive LTD | Al Bahar Al Azraq Drug store | United Kingdom | Diagnostics | Diagnostics | Conditionally Approved as a rapid, near-patient genetic testing system for detection of the MT-RNR1 m.1555A>G variant in newborns considered for aminoglycoside antibiotic therapy. The Genedrive MT-RNR1 ID Kit is CE-marked and intended to guide antibiotic selection to reduce the risk of aminoglycoside-induced hearing loss. Used at point-of-care for rapid (approx. 26 min) detection of the m.1555A>G variant in neonates prior to aminoglycoside administration. Requires buccal (cheek) swab sample collected and processed according to manufacturer’s instructions. Positive results must be confirmed by accredited laboratory genetic testing prior to permanent entry into the medical record. Test operation and result interpretation to be performed by trained healthcare professionals with oversight from clinical genetics or neonatology teams. Implementation must not delay antibiotic administration; antibiotics should be administered within 1 hour of decision to treat. Maintain strict protocols for data privacy and secure handling of genetic and patient data. Updates to test protocols or parameters should follow a documented, stepwise approach with clinical governance oversight. Contionally Approved as per AAMEN: Any health information to be wiped before taking the device out of the facility for maintenance. Population / Intended User: Newborns (neonates) considered for aminoglycoside antibiotic therapy due to suspected bacterial infection To Be Performed By: Trained healthcare professionals in: Neonatology Clinical genetics Paediatric infectious diseases Clinical Setting: Neonatal units, maternity hospitals, or equivalent acute care settings Condition of Use: Used at point-of-care to inform antibiotic selection and reduce risk of aminoglycoside-induced hearing loss in newborns
2025-12-26 | HTA25095 | TFScan | Amplifai health | ARJ Technomed Trading LLC | Saudi Arabia | System/ Software/ AI / App | Screening | Approved as adjunct tool to be adopted Recommendations: 1.TFScan is recommended for use in adult patients (≥18 years) with diabetes as a supplementary tool for the detection and monitoring of diabetic foot ulcers (DFUs) where clinically indicated. 2.To be used by trained and competent healthcare professionals only, in clinical settings. 3.TFScan should not be used as a stand-alone diagnostic tool and must not replace physician assessment, clinical examination, or diagnostic decision-making. 4.Thermal image acquisition must be performed using compatible thermal camera models and in accordance with manufacturer instructions to ensure reliable and interpretable results. 5.Standardized imaging conditions should be ensured, including proper foot positioning, adequate acclimatization time, and minimization of motion artifacts. 6.TFScan should be integrated into diabetic foot screening and follow-up protocols, alongside visual inspection, neurological assessment, and vascular evaluation. 7.Abnormal thermal findings should prompt clinical review, further diagnostic evaluation, or referral according to local diabetic foot care pathways. 8.User training and periodic competency validation should be ensured, including correct image capture, interpretation of outputs, and understanding of software limitations. 9.Clinical outcomes and safety events (e.g., ulcer development, progression, referral rates) should be monitored and documented as part of routine quality assurance and post-market surveillance activity. 10. Company to continuously monitor patients outcomes to ensure technology’s effectiveness and safety and report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division in DOH as part of a Real World Data. Clinical Use Overview: Population / Intended User: Adults (18 years older) with diabetes at risk of DFU To be Performed By: Trained healthcare professionals Clinical Setting: Health care facilities, Hospitals, Clinics and health centers Condition of Use: Screening tool only, results must be interpreted alongside clinical exam findings. Contraindications: TFScan is not intended for use in: 1. Patients under 18 years of age. 2. Non-diabetic foot conditions, including traumatic injuries, acute infections unrelated to diabetes, or vascular conditions not associated with diabetic foot ulcers (DFUs). 3. Open wounds requiring immediate surgical or emergency intervention, where standard clinical assessment should take priority. 4. Situations where thermal images cannot be reliably captured, such as severe motion artifacts, improper camera positioning, or incompatible thermal camera models. 5. Use as a stand-alone diagnostic tool or as a replacement for physician diagnosis or clinical judgment.
2025-12-26 | HTA25093 | NMBA Monitoring & Nerve Location (NMS450X) | Emirates & World Medical Supplies | Xavant Technology | South Africa | Medical Device | Diagnostics | Approved to adopt the technology Recommendations: 1. Recommend the NMS450X Peripheral Nerve Stimulator for hospital use in operating rooms for quantitative neuromuscular blockade (NMBA) monitoring and peripheral nerve mapping/locating, where clinically indicated. 2. Use should be restricted to trained and competent anesthesia professionals only, with documented training on device operation. 3. Use should be integrated into standard anesthesia protocols when neuromuscular blocking agents are administered, including objective documentation of neuromuscular recovery prior to extubating. 4. Strict adherence to manufacturer instructions for use (IFU) is required, including contraindications, warnings related to implanted electrical devices, electromagnetic interference, electrode placement, and infection control. 5. Nerve locating using invasive needles should only be performed by experienced clinicians and avoided in patients with infection at the puncture site or significant coagulation disorders. 6. Clinical outcomes and safety events (e.g. residual neuromuscular blockade, postoperative respiratory events, skin reactions, device-related incidents) should be monitored and documented as part of routine quality assurance and post-market surveillance. 7. The device should not be used as a standalone diagnostic tool for neuromuscular or neurological disorders and should not replace clinical judgment. 8. Initial and periodic staff training and competency validation should be ensured as part of routine clinical governance. 9. Clinical outcomes and device-related safety events should be monitored and documented as part of routine quality assurance and post-market surveillance activities. Clinical Use Overview: Population: Patients undergoing GA with NMBA use; patients undergoing regional anesthesia where mapping/locating is required. To be Performed By: Neurologist, surgeons and qualified anesthesia clinicians (and trained perioperative staff under clinical governance. Clinical Setting: Hospital - Anesthesia (OR-ICU) Condition of Use: Tracking of drug dosage given during the surgeries, tracking of residual at the end of the surgery, tracking of the sedation level in ICU. Contraindications: The NMS450X must NOT be used in patients with: 1. Local infection at the stimulation or needle insertion site– Risk of spreading infection or tissue injury. 2. Known neurological or neuromuscular disorders– Examples: myopathies, peripheral neuropathies, motor neuron disease– Results may be unreliable or misleading. 3. Severe coagulation disorders (especially for invasive nerve locating)– Increased risk of bleeding or hematoma formation. 4. Implanted active electrical devices (relative contraindication)– Pacemakers, ICDs, neurostimulators– Use is not allowed unless cleared by cardiology/biomedical engineering due to interference risk. 5. Neonates (for EMG monitoring)– Device EMG function is not indicated for neonates per IFU limitations.
2025-12-26 | HTA25092 | Lokomat Pro | Amana Healthcare | Hocoma AG (a DIH Group company) | Swizerland | Robotic-assisted gait orthosis with body weight support | Therapeutic | Approved -The Lokomat is a robotic device intended to enable intensive rehabilitative gait therapy in adult and pediatric patients with severe to moderate impairments in walking abilities and functional mobility. • Adoption: Consider acquiring Lokomat for rehabilitation centers with high volumes of neurological gait patients (stroke units, pediatric neuro rehab, SCI centers). Use it as an adjunct to, not replacement for, conventional physiotherapy therapy. Ensure protocols are in place to select suitable patients (e.g. ambulatory potential, weight limits) and to integrate Lokomat sessions with other therapies. • Training & Use: Only trained therapists/engineers should operate the device. Provide formal training on setup (proper harness fitting, strap tension, weight support calibration) and emergency procedures (stop buttons, patient release) as per manufacturer guidelines. Start sessions with safety checks and gradually increase intensity. • Verification & Calibration: Regularly verify device calibration, especially body-weight support accuracy and force sensor function. Implement routine maintenance per the datasheet. • Quality Monitoring: Collect and review clinical outcomes and any adverse events in a rehab registry. Audit that gait metrics (speed, distance). Monitor patient satisfaction and adjust therapy plans as needed. • Patient Safety: Enforce continuous patient monitoring during sessions. Inspect skin under harness/straps regularly for signs of irritation. Adjust support parameters to avoid joint overextension. Keep emergency stop within reach and instruct staff to watch for unusual device behavior. •Install an additional safety mechanism in the Lokomat (Emergency Stop) to safeguard the patient and prevent any incidents. • Population The Lokomat can be used for patients with the following conditions in particular: • Stroke • Traumatic brain injury (TBI) • Paraplegia (SCI) • Cerebral palsy (CP) • Multiple sclerosis (MS) • Parkinson’s disease (PD) • Endoprosthesis (e.g., joint replacements) • Degenerative joint diseases of the lower limbs (e.g. knee osteoarthritis) • Spinal muscular atrophy • Muscle weakness due to lack of mobility • Cardiac conditions if permitted by treating physician • Amputees • Performed by: Trained physiotherapists or rehabilitation clinicians. • Clinical Setting: Healthcare facilities. • Condition of Use: Used as part of a structured gait-training program on the Lokomat treadmill with body-weight support; patients are harnessed and instructed to actively attempt stepping under therapist supervision.
2025-12-26 | HTA25088 | EasyNAT® GBS | Babirus Medical Equipment Trading LLC | Ustar Biotechnologies (Hangzhou) Ltd. | China | Diagnostic Test | Diagnostics | Approved : This kit is an automated, in vitro nucleic acid amplification test used for qualitative diagnosis of group B Streptococcus (GBS) deoxyribonucleic acid (DNA) in vaginal swab. This kit is indicated as an aid in determining GBS colonization status in antepartum and intrapartum women. This kit is not for near patient and self-testing. • Adoption: recommended for healthcare facilities (especially maternity hospitals and birthing centers) consider adopting the EasyNAT GBS Assay as part of their GBS screening protocol. This would be particularly beneficial for intrapartum screening of women who arrive with unknown GBS status, to ensure that those who are colonized receive prompt antibiotic prophylaxis. • Training: Laboratory staff and/or point-of-care operators should be trained on the EasyNAT system, including sample collection technique (proper vaginal swab sampling), cartridge handling, and instrument operation. • Quality Assurance: Implement robust quality control measures. Use the provided positive and negative control materials with each new lot or as required to ensure the test and device are functioning properly . Regularly maintain and calibrate the EasyNAT instrument according to the manufacturer’s schedule. • Monitoring and Evaluation: After implementation, continuously monitor outcomes and performance. Track the rate of neonatal early-onset GBS infection before and after adopting the assay as a quality indicator. Also track the proportion of women receiving antibiotics unnecessarily (false positives or risk-based over-treatment) – the goal is to see that targeted testing reduces unnecessary antibiotic exposure while protecting newborns. Collect feedback from end-users (midwives, lab techs, obstetricians) on the test’s ease of use and impact on workflow. Population: Pregnant women (35–37 weeks gestation or in labor) requiring GBS screening; To be performed by: laboratory technologists or trained maternity ward staff (e.g. midwives) using the EasyNAT system; Clinical setting: hospital labor & delivery unit or point-of-care lab near the bedside; Condition of use: utilized as a cartridge-based rapid NAT at the point of care, following manufacturer’s instructions, to guide intrapartum antibiotic prophylaxis decisions in real time.
2025-12-26 | HTA25062 | Rayvolve | ARJ Medical Supplies LLC | AZMed | France | Medical Device | Diagnostics | Conditionally Approved – Computer-aided diagnosis tool, intended to help radiologists and emergency physicans to detect and localize abnormalities on standard X-rays. Use as AI-assisted decision support tool for X-ray interpretation. Do not use as a standalone diagnostic tool. Not a substitute for clinical judgment. Ensure adequate user training and awareness of limitations. Conditionally approved as per AAMEN team: The approval excludes access to systems from outside UAE. Population: adults and pediatric population. To be performed by: Trained healthcare professionals and radiologists Clinical setting: Health care facilities, radiology, imaging departments Condition of use: Do not use Rayvolve to detect foreign matters. Rayvolve works optimally on chest radiographs taken under optimal conditions: when the patient is facing frontally and in full inspiration.
2025-12-17 | HTA25057 | Omnipod® 5 Automated Insulin Delivery System | ALPHAMED DRUG STORE/Insulet Middle East & Developing Markets as a branch office of Insulet Netherlands B.V | Insulet Corporation | USA | Medical Device | Therapeutic | Conditional Approval of the The Omnipod 5 AID insulin delivery System. It is a tubeless, wearable insulin pump that automatically adjusts insulin every 5 minutes based on continuous glucose monitor (CGM) readings . Population/ Intended User: • Type 1 diabetes aged 2 years and older, • Type 2 diabetes aged 18 years and older. To be performed by: • Patients • Clinical support may assist in setup and monitoring Clinical Setting: Home setting Condition of use: • Must be used with compatible interoperable CGMS • Must follow prescribe insulin settings and regularly monitor for alerts Conditional approval by Information Security AAMEN Team: Mobile application can not be used until clearance from DOH Info Sec is granted. Requestor to revert with the progress status of Planned Control Measures for extension or closure of Exemption. • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2025-12-17 | HTA25104 | MyoBlate RFA | Al Rowais Medical Equipment | RF Medical Company Ltd | South Korea | Medical Device | Therapeutic | Approved as a minimally invasive, uterus-preserving radiofrequency ablation (RFA) system for the treatment of symptomatic uterine fibroids, provided that strict patient-selection criteria, standardized imaging, and safety protocols are followed. 1. Used for women with symptomatic uterine fibroids ≤5 cm, confirmed by ultrasound ± MRI. 2. Requires high-quality imaging to confirm fibroid size, number, type (FIGO 0–5), accessibility, and exclusion of malignancy. 3. Procedure must be performed by gynecologists trained specifically on MYGEN 1000 MYOBLATE RFA technique, ultrasound-guided targeting, and emergency procedures. 4. A structured training pathway must be completed, including device operation, safety features, lesion planning, energy-delivery parameters, troubleshooting, and simulation-based competency. 5. Device use requires strict adherence to clinical protocols for patient selection, informed consent, peri-operative safety, and standardized post-procedure follow-up. 6. All cleaning, disinfection, consumable handling, and contamination-prevention procedures must comply with infection-control SOPs and manufacturer guidance. 7. Any software updates or parameter modifications must follow a documented, controlled process to ensure patient safety and data integrity. Population / Intended User: Adult women with confirmed uterine fibroids causing moderate to severe symptoms who desire a uterus-preserving, minimally invasive treatment. To Be Performed By: Gynecologists consultant or surgeons Clinical Setting: 1. Day surgery units 2. Hospitals with emergency surgical support and ultrasound capabilities Condition of Use: Used for ultrasound-guided RFA treatment of symptomatic uterine fibroids (≤5 cm), including FIGO types 0–5, for women seeking: • Symptom improvement • Uterine preservation • Non-surgical alternatives to hysterectomy/myomectomy • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2025-12-17 | HTA25055 | My Genie -DNA based Genetic testing for prognostic and prevention use | Mekoshaa Alami LLC FZ | MapMyGenome | India | In Vitro Diagnostic | Diagnostics | Conditionally Approved as a DNA-based genetic testing service for personalized risk assessment and preventive health screening. The service is performed by a CAP-accredited laboratory (MapMyGenome, India) and is intended to support individualized, predictive, and preventive medicine by identifying genetic risk factors for chronic conditions such as diabetes and cardiovascular disease. Used as a first-line wellness screening tool for asymptomatic individuals to inform health promotion and early intervention strategies. Testing analyzes over 160 genetic markers to provide personalized insights for disease risk, nutrition, and metabolic health. Results are intended to complement, not replace, standard diagnostic procedures and clinical judgment. All sample collection and data handling must comply with UAE data privacy regulations and ensure secure transfer to the external CAP-accredited laboratory. Genetic counseling and interpretation of results should be provided by qualified healthcare professionals. Any updates to testing protocols or reporting must follow a documented, stepwise approach and be reviewed for compliance with local regulatory requirements. Population / Intended User: Adults seeking preventive health screening and risk stratification for non-communicable diseases. To Be Performed By: Qualified healthcare professionals trained in genetic counseling and preventive medicine. Clinical Setting: Healthcare facilities with services in genetic counseling available Condition of Use: Used for preventive health screening and risk assessment to guide lifestyle and intervention strategies in accordance with international guidelines for genomic screening. Information Security Condition : Requestor to revert with the progress status of Planned Control Measures for closure of Exemption. • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2025-12-17 | HTA25054 | Arthrex | Amico Group | Arthrex Inc. | USA | Medical Device | Others | Conditionally Approved as a preoperative planning system for use in total shoulder arthroplasty with Arthrex-compatible implants. The system is FDA-cleared (Class II) and intended to support implant positioning and surgical planning based on patient specific CT imaging. • Used preoperatively for CT-based 3D planning and implant positioning with Arthrex shoulder systems. • Requires a high-quality CT scan following Arthrex imaging protocols. • Plan creation and modifications must be performed by trained Arthrex planners with surgeon review and approval. • Surgeons must follow VIP instrumentation guidelines to ensure correct transfer of the approved plan to surgery. • Maintain strict protocols to ensure data privacy and secure handling of patient imaging data. • Updates or parameter changes should follow a documented, stepwise approach. Population / Intended User: Adult patients undergoing shoulder arthroplasty To Be Performed By: Orthopedic surgeons specializing in: Shoulder arthroplasty, or Upper limb reconstructive surgery Clinical Setting: Hospitals Condition of Use: Used preoperatively for virtual 3D planning and positioning of implants to enhance precision and surgical outcomes. Conditionally Approved as per AAMEN Team: 1.Requestor to initiate with HFLD for Mobile/Web Application review and approval. 2. Right to audit by regulatory body (DOH) 3. Requestor to revert with the progress status of Planned Control Measures for extension or closure of Exemption. • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. • Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). • Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. • Provision of regular updates and reports about the product to DOH upon request.
2025-11-11 | HTA25075 | Xephilio OCT-R1Optical Coherence Tomography | First Global Medical Equipment Co. LLC | CANON | JAPAN | Medical Device | Diagnostics | Conditional Approval: •To be used in ophthalmology clinics and hospital eye units: For in vivo imaging and measurement of the retina, retinal nerve fiber layer, and optic nerve head. It is intended for use as an aid in the diagnosis and management of posterior segment diseases. It is not intended to be used as the sole diagnositc instrument for disease. A Qualified doctor is responsible for the definitive diagnosis. Used in the diagnosis and monitoring of retinal and optic nerve disorders, including: •Age-related macular degeneration •Diabetic retinopathy •Glaucoma •Macular edema and other posterior segment pathologies Population / Intended User: Patients requiring retinal and optic nerve imaging for diagnosis or monitoring of eye diseases. To Be Performed By: Qualified ophthalmologists, optometrists, or trained ophthalmic technicians. Clinical Settings: Hospitals & eye clinics Condition of Use: Used for non-invasive imaging of the retina and anterior segment to support clinical diagnosis and treatment planning. Conditionally approved as per AAMEN Team : 1.Engagement Contract to be in place and addressing all confidentiality requirements. 2.All remote access sessions to be controlled, logged and monitored. 3.Right to audit by regulatory body (DOH). 4.Requestor to revert with the progress status of Planned Control Measures for extension or closure of Exemption. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2025-11-11 | HTA25080 | Transpara | Smart Medical Solutions | ScreenPoint Medical B.V. | The Netherlands | Medical Device | Diagnostics | Conditional Approval: To be used as a concurrent reading aid for physicians interpreting screening full-field digital mammography exams and digital breast tomosynthesis exams from compatible FFDM and DBT systems, to identify regions suspicious for breast cancer and assess their likelihood of malignancy. It will not replace physician clinical judgment. 1. Transpara (2.1.0) to be used as an AI-assisted reading tool in breast imaging centers to enhance cancer detection accuracy and optimize radiologist workflow, ensuring integration with existing PACS and RIS systems. 2. Develop SOPs for AI deployment, covering image data input/output handling, DICOM compliance, cybersecurity validation, algorithm performance monitoring, and result interpretation workflow. 3. Conduct structured training for radiologists and technologists on Transpara’s functionality, interpretation of region and case scores, and limitations to prevent overreliance on AI outputs. 4. Implement regular performance audits using real-world image datasets to verify sensitivity, specificity, and false-positive rates against benchmark data. 5. Integrate IT and biomedical engineering teams for system calibration, software maintenance, version control, and feedback reporting to the vendor. 6. Enforce data security measures by ensuring encrypted image transfer, controlled access, and adherence to UAE data-residency and patient-privacy regulations. 7. Establish a quality assurance plan with periodic algorithm validation, comparison against radiologist reads, and documentation of any performance drift or anomalies. 8. Maintain continuous vendor collaboration for software updates, patch management, and review of post-market surveillance data to address potential safety or accuracy issues. 9. Promote multidisciplinary oversight (clinical, IT, and QA teams) to review AI impact on workflow, clinical decisions, and patient outcomes. Population / Intended User: Women who meet existing DOH breast-screening eligibility To Be Performed By: Radiologist Clinical Setting: Accredited breast-screening units Condition of Use: To be used as a concurrent reading aid for physicians interpreting screening full-field digital mammography exams and digital breast tomosynthesis exams from compatible FFDM and DBT systems, to identify regions suspicious for breast cancer and assess their likelihood of malignancy. It will not replace physician clinical judgment. Conditionally approved as per AAMEN team: Requestor to revert with the progress status of Planned Control Measures for extension or closure of Exemption. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2025-11-11 | HTA25085 | Tango Reflex Neo™YAGSLT Laser system | First Global Medical Equipment Co. LLC | LUMIBIRD Medical | France | Medical Device | Therapeutic | Conditionally approve as : 1. Selective Laser trabeculoplasty (SLT) Tango Neo, Tango Reflex Neo (SLT mode) for primary open-angle glaucoma (POAG) and ocular hypertension (OHT) 2. Tango Neo, Tango Reflex Neo, UltraQ Reflex Neo (YAG mode): •Iridotomy and iridectomy (angle-closure or prophylaxis) •Posterior capsulotomy (post-cataract posterior capsular opacification – PCO) •Posterior membranectomy Tango Reflex Neo, UltraQ Reflex Neo: Vitreolysis • Adopt for ophthalmology centers with trained laser users and regular equipment maintenance. • Monitor IOP after treatment and record laser settings for each case. • Collect data on IOP reduction and complications for local quality review. Conditionally approved as per AAMEN Team : 1.Engagement Contract to be in place and addressing all confidentiality requirements. 2.All remote access sessions to be controlled, logged and monitored. 3.Right to audit by regulatory body (DOH). 4.Requestor to revert with the progress status of Planned Control Measures for extension or closure of Exemption. • Population: Adults with PCO or glaucoma (OAG/OHT). in glaucoma and cataract services, • Performed by: Ophthalmologists • Setting: Eye clinic or hospital • Conditions of use: Use manufacturer’s guidelines; posterior offset for YAG; standard SLT protocol; monitor post-treatment IOP. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2025-11-11 | HTA25081 | NeuroPace RNS System Programmer | Cleveland Clinic Abu Dhabi | NeuroPace, Inc. | United States of America (USA) | Medical Device | Therapeutic | Conditionally Approved as an adjunctive therapy in reducing the frequency of seizures in individuals 18 years of age or older with partial onset seizures who have undergone diagnostic testing that localized no more than 2 epileptogenic foci, are refractory to two or more antiepileptic medications, and currently have frequent and disabling seizures (motor partial seizures, complex partial seizures and/ or secondarily generalized seizures). The RNS® System has demonstrated safety and effectiveness in patients who average 3 or more disabling seizures per month over the three most recent months (with no month with fewer than two seizures) and has not been evaluated in patients with less frequent seizures. The programmer is not the therapeutic device itself; it is the clinical control point that makes RNS a personalized therapy (detecting each patient’s seizure patterns and tuning stimulation). Population/ Intended User: individuals 18 years of age or older with drug-resistant focal epilepsy (focal onset seizures) who are implanted with the RNS System. To be performed by: Clinicians (neurologists, epileptologists, neurosurgeons) trained in RNS programming. Clinical Setting: Hospitals Condition of use: Used after RNS implantation to interrogate the device, download brain data, adjust detection and stimulation settings, and upload changes. Communication is via the external wand placed over the implant. •Recommended for patients with drug-resistant focal epilepsy •Should be used only in comprehensive epilepsy centers with trained clinicians. •Programming and follow-up must be performed by certified clinicians familiar with the RNS system. •Ensure regular follow up visits (every 1-3 months initially, then as needed) for data download and device adjustment. •Maitian strict protocols to prevent data loss or communication errors between the want, programmer, and implant. •Adhere to infection control and surgical follow-up to minimize post-implant complications. •Clinicians must undergo specialized training by the manufacturer. •Programmer use is restricted to clinical environments. •Updates and parameter changes should follow a stepwise approach. •Use the Patient Data Management System (PDMS) to store brain data securely and support longitudinal care. Conditionally Approved as per AAMEN Team: 1.Requestor to resubmit the application for any change in the data flow or data access. 2.Right to audit by regulatory body (DOH). •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2025-11-11 | HTA25091 | Elecsys Plasma pTau181 | Pharmatrade LLC | Roche Diagnostics | Germany | Diagnostic Test | Diagnostics | Approved as an in vitro quantitative immunoassay intended for the determination of the phosphorylated Tau 181 protein in human plasma. The Elecsys Phospho Tau (181P) Plasma assay result is intended to be used for the early detection of amyloid pathology in adult subjects aged 55+ years with subjective complaints or non-specific cognitive impairment (subjective cognitive decline) or whom their physician suspects of having mild cognitive impairment and who are being evaluated for Alzheimer's disease and other causes of cognitive decline in primary, secondary, or tertiary care settings. It is emphasized that the assay is an adjunct tool rather than a standalone diagnostic. • Adopt Test in Appropriate Patients: Healthcare providers and systems should consider implementing Elecsys pTau181 testing for adults 55+ with unexplained mild cognitive impairment or early dementia, in order to triage for Alzheimer’s disease pathology. Using this blood test at the initial assessment stage (particularly in primary care or general neurology clinics) can help identify patients unlikely to have AD, thus streamlining referrals. • Ensure Laboratory Verification and Training: Before adopting the assay, laboratories should perform verification of test performance and cutoffs in their own setting. This includes running precision studies, cross-checking calibration, and confirming the manufacturer’s reference range or cutoff with local sample populations . • Quality and Regulatory Compliance: Maintain rigorous quality assurance practices when utilizing this test. • Monitor Emerging Evidence and Update Practices: As real-world data accumulate and new studies on blood AD biomarkers are published, healthcare providers should stay updated on the latest guidelines and performance data. Population: Adults 55–80 with cognitive impairment symptoms To be performed by: Clinical laboratory professionals Clinical setting: Healthcare facility Condition of use: In vitro diagnostic assay, one-time plasma test as an adjunct to Alzheimer’s evaluation (used to rule out AD amyloid pathology, with results integrated into overall diagnostic workup). Approved as per AAMEN team considering the following condtions : -PHI/PII shall not be stored, transferred, or made available outside UAE -Support staff (L1, L2 and L3) shall have no visibility of PII. -L1 and L2 support shall be within UAE.
2025-10-29 | HTA25063 | Velys robotic-assisted solution (Attune® knee system ) | Pharmatrade / Johnson & Johnson (Middle East) Inc. | DePuy Ireland UC (a Johnson & Johnson company) | Ireland / USA (manufacturing sites include DePuy Ireland UC – Loughbeg, Ringaskiddy, Co. Cork, Ireland; Viant Brimfield, Viant Fort Wayne, Anspach USA, Synthes USA, RMS Surgical Canada, Symmetry Manchester UK, Komet Germany) | Medical Device | Therapeutic | Approve: 1. Define patient selection criteria focusing on primary total knee arthroplasty (TKA) candidates with osteoarthritis or degenerative knee conditions, while excluding patients with severe bone loss, active infection, or contraindications for robotic intervention. 2. Develop and integrate standard operating procedures (SOPs) covering all stages of robotic-assisted TKA from preoperative imaging and intraoperative robotic calibration to postoperative follow-up and outcome tracking. 3. Establish a structured multidisciplinary training program for orthopedic surgeons, scrub nurses, anesthesiologists, and biomedical engineers. This should include hands-on cadaver labs, simulation modules, proctor-led live cases, and periodic refresher workshops to ensure surgical precision and device safety. 4. Ensure robust biomedical engineering support by setting up protocols for robotic system calibration, predictive maintenance, and troubleshooting, with certified biomedical engineers to reduce downtime. 5. Promote integration with national infection control and safety protocols, emphasizing sterile workflows and adherence to device-specific calibration and sterilization procedures to minimize risks of mechanical malfunction or infection. Criteria for use: Population / Intended User: Adult patients undergoing primary TKA. To be Performed by: Orthopedic surgeons trained and certified in VELYS™ operation. Clinical Setting: Hospital operating theatres with robotic-assisted orthopedic surgery capabilities. Condition of Use: Under sterile surgical conditions, with compatible DePuy Synthes knee implant systems. Contraindications: Patients with conditions precluding TKA or where robotic-assisted intervention is not advised (severe bone loss, active infection, inability to tolerate anesthesia). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2025-10-29 | HTA25051 HTA25061 | Trucheck Intelli | HPI MEDICAL SOLUTIONS LLC SP Mine Gene (FZC) | Datar Cancer Genetics UK Private Limited | United Kingdom | Diagnostic Test | Diagnostics | Conditional Approval Trucheck is, non-invasive blood test that detects circulating tumor cells (CTCs) 1. Trucheck is a screening test, not a diagnostic tool. A positive result does not confirm cancer, and a negative result does not rule out cancer. 2. Further Clinical Evaluation Required:All test results, especially positive ones, must be corroborated by imaging, biopsy, or other standard diagnostic procedures. 3. Trucheck is designed for asymptomatic individuals at risk of cancer. It is not intended for use in individuals with signs or symptoms of cancer. 4. Clinical Judgment Is Essential: Test results should not be used in isolation. Clinical history, other lab tests, and imaging findings must all be considered. 5. Not Suitable for Monitoring or Therapeutic Decisions: The test is not validated for tracking cancer progression, treatment response, or recurrence. 6. Condition by Information Security: Requestor to revert with the progress status of Planned Control Measures for extension or closure of Exemption. Population/ Intended User: Trucheck™ is designed for asymptomatic individuals who are at an elevated risk of developing solid organ cancers. Primarily recommended for individuals aged 40 to 70, though it can be considered for younger individuals with high cancer risk; strong family history, hereditary mutation carriers) under specialist supervision. To be performed by: Healthcare professionals Clinical Setting: clinical laboratories Condition of use: early detection screening tool for adults aged 40-70 ( or younger individuals with increased cancer risk) with no prior cancer diagnosis. •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request.
2025-10-29 | HTA25059 | BrainsWay Deep TMS Therapy | Gaelan Medical Trade Single Owner LLC | BRAINSWAY LTD | ISRAEL | Medical Device | Therapeutic | • BrainsWay Deep TMS™ is conditionally approved for adoption for the following indications in adult patients suffering from neuropsychiatric or neurological disorders, chronic pain and smoking cessation who have not achieved satisfactory improvement with standard treatments: 1. Adults with Depression – MDD/TRD 2. Adults with Anxiety comorbid with MDD 3. Adults with OCD 4. Adults with PTSD in refractory cases with criteria for use 5. Adults with PSUD/Addiction – Alcohol with criteria for use 6. Adults with Chronic neuropathic Pain as an adjunctive treatment with criteria for use 7. Adults with Smoking Cessation • Conditionally approved by DOH INFOSec: Requestor to revert with the progress status of Planned Control Measures for closure of Exemption. • The system should be used as an adjunct therapy in patients already receiving standard of care treatments, including pharmacotherapy and/or psychotherapy, in accordance with the Instructions for Use (IFU). • The device must be operated by trained psychiatrists, neurologists, or qualified healthcare professionals who are certified in Transcranial Magnetic Stimulation protocols. • Patient eligibility screening must be conducted before each treatment course to identify contraindications such as seizure risk, metallic cranial implants, or other conditions listed in the IFU. • Treatment protocols (frequency, intensity, number of pulses, and session duration) should be individualized based on the clinical indication and established evidence-based TMS protocols. • Continuous patient monitoring during stimulation sessions is recommended to detect adverse reactions such as headaches, scalp discomfort, or rare seizure events. • The technology should be integrated within multidisciplinary mental health treatment pathways, including psychiatry, neurology, addiction medicine, and behavioral health services. • Clinical outcomes and safety data (e.g., depression scales, OCD severity scores, smoking cessation rates, pain scores, PTSD symptom scales) should be systematically recorded as part of post-market surveillance and real-world evidence generation. • Facilities implementing Deep TMS should ensure periodic staff training, competency validation, and adherence to regulatory and safety protocols. Clinical Use Overview: Population / Intended User: Adult patients suffering from neuropsychiatric or neurological disorders who have not achieved satisfactory improvement with standard treatments. Indications: • Adults with Depression – MDD/TRD • Adults with OCD • Adults with Smoking Cessation • Adults with Anxiety comorbid with MDD • Adults with PTSD with criteria o DSM‑5 diagnosis of PTSD o Failure of ≥1 evidence‑based psychotherapy (e.g., trauma‑focused CBT, EMDR) o Failure or intolerance of ≥1 pharmacological therapy o Use as adjunct, not monotherapy o Excludes acute PTSD or early intervention • Adults with Chronic neuropathic Pain as an adjunctive treatment o Adults ≥18 years o Diagnosis of chronic neuropathic pain (duration ≥3–6 months) o Examples include: o Diabetic peripheral neuropathy o Chronic peripheral neuropathic pain of non acute origin o Documented failure of standard neuropathic pain treatments (e.g., anticonvulsants, antidepressants) o Adjunctive use only (not monotherapy) Explicit exclusions: o Non neuropathic chronic pain o Inflammatory pain disorders o Migraine and headache syndromes • Adults with PSUD/Addiction – Alcohol o Adults ≥18 years o Diagnosis of moderate–severe Alcohol Use Disorder o Failure of standard psychosocial and/or pharmacological treatments o Delivered only as part of a comprehensive addiction program o Device‑restricted to BrainsWay (H7 coil targeting mPFC/ACC) o Not approved for primary prevention or mild AUD To Be Performed By: Psychiatrists, neurologists, trained TMS clinicians, and certified healthcare professionals. Clinical Setting • Hospitals • Specialized mental health centers • Neurology clinics • Outpatient psychiatric clinics Condition of Use: Non-sterile clinical environment; intermittent outpatient therapy sessions typically delivered over multiple treatment sessions (e.g., daily sessions for several weeks). Contraindications 1. Implanted metallic or electronic devices in or near the head (e.g., cochlear implants, deep brain stimulators, intracranial electrodes). 2. History of epilepsy or seizure disorders, unless evaluated and cleared by a specialist physician. 3. Presence of ferromagnetic material in the cranial region. 4. Severe neurological conditions where cortical stimulation may pose additional risk. 5. Unstable psychiatric conditions requiring urgent medical intervention. 6. Pregnancy.
2025-10-21 | HTA25065 | PulseSelect | Al Naghi Medical Company | Medtronic | Malaysia | Medical Device | Therapeutic | •Approve Pulse Select™ Pulsed Field Ablation (PFA) System electrophysiology centers through a structured, phased rollout under proctored supervision to ensure safe, standardized, and effective introduction of PFA technology for atrial fibrillation (AF) ablation procedures. •Define patient selection criteria focusing on paroxysmal and persistent AF patients with stable anticoagulation, no active infection, and no contraindications to vascular access or general anesthesia. Exclude patients with left atrial thrombus, uncontrolled heart failure, or recent stroke/TIA, and ensure adequate pre-procedural ima ging, laboratory assessment, and anesthesia clearance. •Develop comprehensive Standard Operating Procedures (SOPs) covering all stages of the procedure—pre-procedural preparation, sterile setup, generator calibration, phrenic nerve and esophageal monitoring, mapping strategy, ablation energy parameters, emergency response, and post-procedure observation—to ensure consistency, safety, and traceability. •Implement strict single-use and contamination-control measures for all catheters, sheaths, and accessories. Maintain full traceability of consumables through lot documentation and ensure compliant disposal per infection control and biomedical waste management standards. •Establish a structured training and credentialing program for electrophysiologists, nurses, and biomedical engineers, including hands-on simulation, supervised initial cases, and periodic competency reassessment to maintain operator proficiency. •Embed biomedical engineering support into the operational pathway to conduct acceptance testing, regular calibration, firmware updates, and preventive maintenance per technical service schedule. Implement immediate device quarantine and root-cause analysis in case of malfunction or procedural incident. Population / Intended User: Patients with atrial fibrillation (paroxysmal or persistent) who are candidates for catheter-based pulmonary vein isolation (PVI) and are refractory or intolerant to antiarrhythmic drug therapy. To Be performed by: Healthcare Professionals Clinical settings: Hospitals Condition of use: Used as a single-use, sterile catheter in conjunction with the PulseSelect™ PFA system for performing cardiac tissue ablation. Contraindications: The PulseSelect™ PFA loop catheter is contraindicated for use in patients with the following conditions: •Active systemic infections •A known sensitivity to Heparin •Blood clotting abnormalities •Permanently implanted metallic objects in the left atrium The catheter is also contraindicated in conditions where the manipulation of the catheter within the heart would be unsafe, such as intracardiac mural thrombus. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-10-21 | HTA25069 | Occlutech Atrial Flow Regulator | ELITE CARE MEDICAL SUPPLIES LLC | Occlutech Tıbbi Ürünler San. Ve Tic. Ltd | Turkey | Medical Device | Therapeutic | •Conditioal approval of Atrial Flow Regulator (AFR) only within specialized cardiac centers performing transseptal catheterization procedures for symptomatic heart failure patients (HFrEF or HFpEF) unresponsive to optimal medical therapy. •Conditional to be restricted to for use within a registry to document Real World Evidence outcomes ( NYHA functional class change from baseline, 6MWT and left atrial pressure LAP) and Adverse Events. This is to be done in alignment with medical research center for outcomes and safety documentation that is expected to be reported at end of 12 months from this conditional approval, and till 24motnhs). •Define clear patient-selection criteria, including documented hemodynamic confirmation of raised left atrial pressure and exclusion of right-heart dysfunction or pulmonary vascular disease, ensuring AFR implantation is reserved for eligible high-risk patients. •Develop a structured training program for interventional cardiologists, cardiac anesthetists, and cath-lab nurses focusing on AFR implantation techniques, hemodynamic assessment, and management of potential complications (e.g., embolization, bleeding, arrhythmia). •Integrate AFR follow-up protocols with hospital electronic systems to enable scheduled echocardiographic evaluations, anticoagulation tracking, and automated adverse-event reporting to the device registry. •Implement post-market clinical follow-up (PMCF) and establish a local patient registry capturing safety, device patency, NYHA class, 6-minute walk distance, and hospitalization data, in alignment with EU MDR Class III device requirements. •Embed risk-mitigation measures such as immediate access to surgical backup for device retrieval, dual-antiplatelet therapy for 3–6 months, and defined management pathways for device migration, thrombus, or infection. •Reports to be submitted to DoH PV or real-world data team for long term use and if there are concerns or safety issues. Population / Intended User: Heart failure patients (HFrEF or HFpEF) with elevated left atrial pressure despite optimal medical therapy. To Be Performed By: Intervnetional cardiologists Clinical Settings: Specialized cardiac Hospitals Condition of Use: Implanted to create a controlled left-to-right atrial shunt under imaging guidance to reduce left atrial pressure. •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-10-21 | HTA25074 | MSK-IMPACT | National Reference Laboratory, M42 | Sophia Genetics | USA | Diagnostic Test | Diagnostics | Conditionally approve MSK-IMPACT as a next generation sequencing test. The test to be used only at the NRL where it was validated. The test analyzes hundreds of cancer-related genes in a patient’s tumor sample to identify mutations, copy number changes, and select fusions that may guide diagnosis, prognosis, or treatment. Population/ Intended User: Patients with solid tumors (including metastatic/ advanced cancer) To be performed by: Trained laboratory professionals. Clinical Setting: Healthcare facilities Approved only at NRL labs. Condition of use: Tumor genomic profiling to guide diagnosis, prognosis, and treatment planning. •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-10-21 | HTA25073 | Fuji Lumipulse pTau217 -NfL plasma assays | National Reference Laboratory LLC | Fujrebio | Japan | Diagnostic Test | Diagnostics | Conditionally approve Fuji Lumipulse pTau217 -NfL plasma assays. The test to be used only at the NRL where it was validated. The Lumipulse® G Neurofilament Light Chain (NfL) Plasma Assay is an in vitro diagnostic test that quantitatively measures NfL in human plasma using the Lumipulse G system. Under Research supervision submission of regular reports to HTA and Medical Research teams at DOH. Population / Intended User: Adults with suspected or confirmed neurodegenerative or neuroinflammatory disorders. To be performed by: Trained healthcare professionals. Clinical settings: Clinical and hospital laboratories. Condition of use: Aid in monitoring neuroaxonal injury associated with neurological diseases (not a stand-alone diagnostic test). •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-10-21 | HTA25040 | Accu-Chek SmartGuide Device with SmartGuide App | PHARMATRADE LLC | Roche Diabetes Care GmBH, Sandhofer Strasse 116 68305 Mannheim, Germany | Romania | Medical Device | Diagnostics | Conditionally approve: 1. Accu-Chek SmartGuide Device with SmartGuide App Continuous Glucose Monitoring (CGM) System (a real time, continuous glucose monitoring device) indicated for 2. Applicant to revert with the progress status of Planned Control Measures for closure or extension of Exemption to DOH Information Security 3. Right to audit by regulatory body (DOH). 4. Develop a structured training curriculum Including hands-on training for patients on how to apply the sensor, and provide guidance to physicians for remote monitoring and patient follow-up. Population/ Intended User: – Adults, 18 years of age and older with diabetes mellitus, especially those whoh ave type 2 diabetes on insulin, at risk of hypoglycemia or with type 1 diabetes. To be performed by: – Healthcare Professionals/ Patinets. Clinical Setting: – Healthcare facilities / Patients level . Conditions for use - Prescription only. Contraindications: •The CGM device shall not be used by critically ill patients or patients on dialysis. •The sensor must be removed prior to entering medical treatment areas with high-powered medical electric equipment e.g., magnetic resonance imaging (MRI), computed tomography (CT), or X-ray, radiotherapy, diathermia. •The app is not intended to control any functions or modify any data of the connected devices. •The app is not intended for calculation of insulin doses or other medications. •The app is not intended to replace the primary display of the connected devices (like BG meter, insulin pens or CGM data on the respective CGM app). •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-10-21 | HTA25086 | Abbott i-STAT PT plus cartridge | Abbott Laboratories GMBH | Abbot Point of Care | USA | Diagnostic Test | Diagnostics | Approve "Abbott i-STAT PT plus cartridge" for checking PT/INR in patients taking warfarin. •Make sure staff are trained and perform regular quality checks on the device. •Keep a record of all test results, cartridge lots, and operator IDs for safety and traceability. Population / Intended User: Patients taking warfarin (coumarin) anticoagulant therapy who need regular PT/INR monitoring to manage blood-thinning levels. The device is only for warfarin patients, not for those taking newer blood thinners (DOACs like Eliquis or Xarelto). •If the INR is 4.5 or higher, or the result looks unusual, repeat or confirm in the lab. To Be Performed By: Trained healthcare professionals such as nurses, laboratory staff, or clinicians familiar with point-of-care testing and INR interpretation. Clinical Setting: Hospitals, emergency departments, outpatient clinics, and anticoagulation (warfarin) clinics where rapid coagulation results are needed. Condition of Use: Use fresh whole blood (venous or capillary/fingerstick). Cartridges must be stored at 2–8 °C and used before expiry Note that lupus anticoagulant can slightly affect results.. Follow i-STAT 1 analyzer prompts and run built-in quality control checks before testing. Contraindications: Not for patients using DOACs (Direct Oral Anticoagulants) like Eliquis, Xarelto, Pradaxa, since PT/INR is not useful for those drugs. Avoid use if hematocrit < 24 % or > 54 % (may affect results). Do not use if sample is clotted, hemolyzed, or contaminated. Do not use expired or improperly stored cartridges •Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). •Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. •Provision of regular updates and reports about the product to DOH upon request. •Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. •Compliance with the relevant laws and standards of the Department of Health. •Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-09-29 | HTA25072 | Hepatic Artery Infusion Pump (HAIP) | Burjeel | Intera Oncology | USA | Medical Device | Therapeutic | Approve the Intera 3000 pump only in oncology hospitals with Multidisciplinary tumor board (MDT) approval: Case reviewed and documented (hepatobiliary surgery, interventional radiology, medical oncology, hepatology, anesthesia, nuclear medicine). 1) Define Patient Selection Criteria o Prioritize adults with liver-dominant colorectal metastases (adjuvant or unresectable) or unresectable intrahepatic cholangiocarcinoma. (approve as per the population detailed below in the Criteria for use). o Exclude patients with extensive extrahepatic disease, poor hepatic reserve, active infection, or those unable to comply with regular refills and monitoring. 2) Develop Standard Operating Procedures (SOPs) that covers operational elements: o Pump implantation (surgical protocols, catheter placement, infection prevention). o Pre-treatment mapping and flow confirmation. o Drug sourcing, preparation and refill protocols (aseptic technique, flow verification). o MRI procedures for MR-conditional scanning and post-MRI flow checks. o Toxicity management (especially biliary injury and catheter complications) or any other adverse events related to the machine 3) Structured Training Curriculum: train hepatobiliary surgeons, oncologists, interventional radiologists, nuclear medicine physicians, nurses, biomedical engineers, and pharmacists. Include: o Hands-on sessions on implantation, refill technique, troubleshooting flow issues. o Simulation training for managing complications (e.g., extrahepatic perfusion, infection, pump occlusion). o Periodic refreshers and competency reassessments for all staff. 4) Ensure patients understand the importance of scheduled refills and infection precautions. 5) Provide education on signs of pump malfunction, infection, or hepatic toxicity and establish clear escalation pathways. 6) Enforce strict aseptic refill protocols using non-coring needles, ensuring alignment with DoH infection control standards for implantable devices. 7) Clinical Monitoring Documentation o Baseline and pre-infusion scans (angiography + Tc-99m MAA) to confirm hepatic-only perfusion. o Regular liver function tests (LFTs) to detect early biliary injury or hepatotoxicity. o Residual volume and flow rate checks at every refill visit. o Adverse event tracking and root-cause analysis for complications (e.g., biliary toxicity, catheter thrombosis, pump infection) and must be reported immediately to the PV team in DoH. 8) Monitor and document MRI exposures and perform post-MRI pump checks to maintain safety and compliance. 9) Real-World Data reporting to DOH on annual basis for all operated cases, this should be aligned with Real World Data Section at the Medical research division: MRD@doh.gov.ae a. Clinical Effectiveness • Overall Survival (OS) (months from pump implantation). • Progression-Free Survival (PFS) / Hepatic Progression-Free Survival (hPFS) • Objective Response Rate (ORR) by RECIST (Response Evaluation Criteria in Solid Tumors) and mRECIST (modified RECIST). • Radiological shrinkage of liver tumors, relevant if conversion to resection is an intended outcome. • Conversion to resection/ablation (for mCRC): % of patients converted from unresectable → resectable after HAIP. • Disease Control Rate (the percentage of patients whose cancer shows Complete Response (CR), Partial Response (PR), or Stable Disease (SD). This is to capture partial benefit even if full shrinkage isn’t achieved. 10) Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 11) Provision of regular updates and reports about the product to DOH upon request. General Requirements: • Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. • Compliance with the relevant laws and standards of the Department of Health. • Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting. Please note that this product will be monitored and followed up by the Department of Health. In the event that any complications or side effects are proven as a result of using this product, the Department of Health has the authority to withdraw, suspend and investigate this product in the health sector of the Emirate of Abu Dhabi. Criteria for Use 1) Population/Intended User - Adults with: o Colorectal cancer liver metastases (CRLM) (liver-dominant) — for post-resection (adjuvant) or unresectable liver-dominant disease  The aim is to control and/or conversion to resection/ablation.  Molecular documentation: RAS/BRAF/MSI status recorded to inform systemic sequencing.  Extrahepatic disease if present must be low volume and controlled; systemic plan documented  Prior therapy: intolerance or ineligibility to standard systemic therapy (fluoropyrimidine, oxaliplatin, irinotecan ± biologics; ± immunotherapy if MSI-H/dMMR). o Intrahepatic cholangiocarcinoma (iCCA) liver-confined — for unresectable, liver-dominant cases.  Locoregional positioning: Considered after ineligibility/failure of (or non-response/intolerance to) other locoregionals (e.g., Y-90, TACE, ablation) or when Multidisciplinary tumor board (MDT) documents HAIP as preferable due to tumor distribution/biliary anatomy.  Transplant candidacy: Not a candidate for, or declined, liver transplantation.  Prior therapy: Progressed after or intolerant to gemcitabine-platinum ± immunotherapy/targeted therapy as appropriate. (First-line use may be considered in clinical trial settings or exceptional MDT-documented cases aimed at conversion to resection.) - Patients should have adequate hepatic function, an ECOG performance status of 0–1, and be capable of adhering to scheduled refills and monitoring. - Not intended for pediatric use. 2) To Be Performed By Multidisciplinary tumor board: • HPB/Hepatobiliary Surgeons: For pump implantation and catheterization. • Interventional Radiologists: For pre-infusion angiography, arterial mapping, embolization of extrahepatic branches, and troubleshooting catheter issues. • Medical Oncologists: For designing chemotherapy protocols (e.g., FUDR dosing, integration with systemic therapy), monitoring toxicity, and making dose adjustments. • Oncology Pharmacists: For compounding and verifying HAI chemotherapy to ensure safe refills. • Specialist Infusion Nurses: For performing pump refills with non-coring needles, checking residual volume, and educating patients. • Nuclear Medicine Physicians: For conducting pre-treatment perfusion scans and repeat studies if perfusion changes are suspected. 3) Clinical Setting - Oncology center or Hospital - Requires access to: o Operating room (for pump implantation). o Interventional radiology suite (for arterial mapping, embolization, catheter troubleshooting). o Oncology infusion clinic. o Nuclear medicine unit. 4) Contraindications Intera 3000 HAI Pump Safety Information for Patients The Intera 3000 Hepatic Artery Infusion Pump is contraindicated for use in patients with - Known or suspected infection, bacteremia, septicemia or peritonitis. - Known allergic reaction or other signs of intolerance to implanted devices. - Emotional or psychiatric problems. - Insufficient body size to accommodate the physical size of the pump. - JND Therapeutics Floxuridine for Injection, USP should be used with added caution in patients with impaired hepatic or renal function. - Patients with known disease extending beyond an area capable of infusion must be considered for systemic therapy with other chemotherapeutic agents. - Contraindications relating to the specific drug to be used must be observed and followed per the approved drug labeling. 5) Safety and Side effect documented and reported to DOH PV team • Treatment-related Grade ≥3 adverse events (AEs) • Biliary sclerosis, cholangitis, hepatic failure, GI ulceration. • Pump-related complications • Catheter thrombosis, dislodgement, infection, extrahepatic perfusion. • Treatment discontinuation rates due to toxicity or complications. • Hospitalizations (frequency, length of stay) related to HAIP.
2025-09-11 | HTA25067 | WILAFLOW ELITE NEONATAL VENTILATOR | HESSA MEDICAL EQUIPMENTS CO LLC | WILAMED GmbH | GERMANY | Medical Device | Therapeutic | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. •Population / Intended User: premature and newborn infants. •To Be performed by: Trained healthcare professionals •Clinical settings: Hospital (NICU) •Condition of use: The ventilator provides gentle respiratory support through modes such as SNIPPV / NIPPV, NCPAP and HFNC modes •WILAflow Elite to be used as a non-invasive ventilation solution in NICUs to improve neonatal respiratory outcomes and reduce reliance on invasive ventilation. •Ensure proper training and maintenance to preserve safety and diagnostic accuracy, it is essential to train users on proper handling, hygiene, and interpretation of results, and to follow routine calibration and maintenance protocols. •Monitor data utilization within the device, ensuring stored patient data is managed securely and complies with privacy and information security regulations within Abu Dhabi. 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-09-11 | HTA25071 | SOPHIE VENTILATOR | HESSA MEDICAL EQUIPMENTS | Fritz Stephan | Germany | Medical Device | Therapeutic | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. •Population / Intended User: preterm infants, infants, and small children up to 25 kg, but especially for the ventilation of immature preterm infants with a very low birth weight. •To Be performed by: authorized and properly trained professional healthcare staff. •Clinical settings: Hospital (NICU) •Condition of use: It is sole responsibility of the user to select the most appropriate type of ventilation based on the patient’s medical condition. The continuous monitoring of the patient’s condition must be assured at all times. Sophie must not be used: •In the vicinity of magnetic resonance imaging systems (MRT, NMR, NMI). •In the potentially explosive environments. In positive pressure (hyperbaric) chambers Non-invasive ventilation (NIV) is contraindicated in the following cases: •Nonspontaneous breathing, •Fixed or functional airway obstruction . •Gastrointestinal bleeding or ileus. 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-09-11 | HTA25064 | INUS IN300; Inuspheresis/ Therapeutic plasma filtration | Yas Clinic Khalifa City | INUS Medical Devices AG (Ayus Medical Devices AG) | Swizerland | Technique/ Procedure | Therapeutic | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. •Population/ Intended User: Adults with lipid, or toxin-related conditions. •To be performed by: Trained physician+ nurse •Clinical Setting: Hospital with clear SOP in place to minimize cross-contamination and infection risks, aligning with national infection control guidelines. •Condition of use: Supportive therapy via venous access (2-3 hrs per session) 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-09-11 | HTA25066 | Fuji Lumipulse G pTau217 /P-Amyloid 1-42 Plasma Ratio | National Reference Laboratory LLC | Fujrebio | Japan | Diagnostic Test | Diagnostics | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. •Population / Intended User: Adults that are more than or equal to 50 years with signs of cognitive decline. •To Be performed by: Trained healthcare professionals •Clinical settings: Clinical and hospital laboratories •Condition of use: Aid in identifying amyloid pathology linked to Alzheimer’s disease ( not a stand-alone diagnostic test) 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-09-11 | HTA25038 | DigiStain Cancer risk assessment interpretive software | Anovo Prime Software Trading LLC | Digistain Ltd | United Kingdom | System/ Software/ AI/ App | Others | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. •Population/ Intended User: – Breast Cancer patients •To be performed by: – oncologists and pathologists. •Clinical Setting: – Healthcare facilities / Labs Conditions for use - Hospital Readiness – Requires integration into oncology workflows. Contraindications – Strict adherence to risk classification guidelines to prevent misinterpretation. •Conditionally approved :Requestor to revert To DOH Info Sec with the progress status of Planned Control Measures for closure of Exemption. 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-09-11 | HTA25068 | Abbott i-STAT High Sensitive Troponin I assay | Abbott Laboratories GMBH | Abbott Point of Care | USA | Diagnostic Test | Diagnostics | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. •Population / Intended User: Patients with suspected acute coronary syndrome (ACS) •To Be performed by: Trained healthcare professionals •Clinical settings: Point of care, emergency departments, critical care and laboratory •Condition of use: To support the early diagnosis of myocardial infarction (MI) and the risk stratification of patients with suspected acute coronary syndrome (ACS), complementing clinical evaluation, ECG, and laboratory findings. 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-08-26 | HTA25043 | Us2.ai Cardia Software | Emerging Global Technologies | Eko.ai Pte Ltd | Singapore | System/Software/ AI/App | Diagnostic | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. 4.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 5.Compliance with the relevant laws and standards of the Department of Health. 6.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-08-26 | HTA25056 | RAPIDPoint® 500e Blood Gas System | ONE HEALTH LLC | Siemens Healthcare Diagnostics | GB | Medical Device | Diagnostics | 1.Any adverse events needs to be reported to PV 2.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 3.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4.Provision of regular updates and reports about the product to DOH upon request. 5.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 6.Compliance with the relevant laws and standards of the Department of Health. 7.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-08-26 | HTA25060 | Physidia | Majestic Medical / Aldar Maison Home Healthcare (Abu Dhabi) | Physidia {Saint Barthelemy-d Anjou-France} Made In France | France | Medical Device | Therapeutic | 1.Physidia S3 Portable Hemodialysis System can be used in Abu Dhabi’s home healthcare and satellite clinic programs to support decentralization of renal care, particularly for stable end-stage kidney disease (ESKD) patients. 2.Define patient selection criteria clearly, prioritizing stable adult patients with vascular access, low interdialytic weight gain, and adequate caregiver support. Exclude unstable patients, those with significant comorbidities, or those lacking a suitable home environment (power, drainage, clean space). 3.Develop Standard Operating Procedures (SOPs) for all stages of treatment—device setup, vascular access management, therapy monitoring, waste disposal, and emergency escalation—integrating them into nephrology service protocols. 4.Establish a structured training curriculum for caregivers, nursing staff, and biomedical engineers. Training should include: •Multi-day hands-on sessions on setup, troubleshooting, and vascular access care. •Simulation-based practice for handling alarms, hypotension, or clotting events. •Periodic refresher and competency reassessments. 5.Implement home suitability assessments before deployment, evaluating the household for adequate power supply, drainage, storage of dialysate bags, and infection control measures. 6.Create clinical monitoring protocols including: •Regular in-person follow-ups for adequacy (weekly Kt/V), nutrition, anemia, and vascular access checks. •Adverse event tracking and root-cause analysis for underdialysis or technical failures. 7.Prioritize infection prevention by enforcing the single-use design (sterile dialysate bags, disposable cartridges) and aligning with UAE infection control standards. Regular audits should verify adherence to disposables’ handling and disposal practices. 8.Integrate biomedical engineering support into care pathways to ensure timely preventive maintenance, periodic calibration, and corrective actions per Physidia’s servicing schedule. 9.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 10.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 11.Provision of regular updates and reports about the product to DOH upon request. 12.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 13.Compliance with the relevant laws and standards of the Department of Health. 14.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-08-26 | HTA25052 | Corsano CardioWatch 287-2 System | Response Plus Medical Services | Corsano | Netherlands | Medical Device | Monitoring | 1.Always direct patient to consult Healthcare Professional and to seek medical advice from qualified healthcare providers regarding any health concerns or before making changes to prescribed treatments. 2.Exclusion Not for High-Acuity Environments: The CardioWatch is intended for use in non-acute settings, such as general hospital wards or at home, and is not suitable for high-acuity environments like intensive care units (ICUs) or operating rooms. 3.Not a Substitute for ECG Monitoring: This device does not replace continuous ECG monitors. Patients with acute cardiac conditions requiring detailed ECG analysis should use appropriate ECG monitoring equipment. 4.Limited Performance in Certain Conditions: The CardioWatch may not provide accurate SpO₂ readings in situations involving high motion or low blood perfusion. 5.Please note that this product will be monitored and followed up by the Department of Health. In the event that any complications or side effects are proven as a result of using this product, the Department of Health has the authority to withdraw, suspend and investigate this product in the health sector of the Emirate of Abu Dhabi. 6.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 7.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 8.Provision of regular updates and reports about the product to DOH upon request. 9.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 10.Compliance with the relevant laws and standards of the Department of Health. 11.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-07-31 | HTA25048 | MammoScape | ARJ Medical Supplies LLC | MeVis BreastCare GmbH & Co. KG 4 | Germany | System/Software/ AI/App | Diagnostic | 1.Adopt MammoScape™ across breast-imaging units that perform screening, diagnostic work-ups, and surgical-planning reviews for breast-conserving surgery. 2.Define clear case-selection criteria, all screening mammograms, digital breast tomosynthesis (DBT), contrast-enhanced mammography, and pre-operative MRI, so only validated image types enter the MammoScape workflow, while emphasizing that 3.AI interpretation does not replace clinical judgment, final decisions must always remain with the attending physician. 4.Create end-to-end standard operating procedures (SOPs) covering image acquisition, DICOM transfer, AI-score display, report generation, and archive/backup, and embed these SOPs into the radiology information system (RIS) and PACS. 5.Develop a structured training curriculum for radiologists, surgeons, medical physicists, and application engineers, including hands-on workstation use, AI-awareness modules, interpretation drills, and routine competency refreshers. 6.Integrate MammoScape with hospital IT that includes scheduled deployment of vendor software-pack updates (AI engine, workstation OS), DICOM-conformance upgrades and quarterly cybersecurity patches; validate each release in a staging environment, document change control and maintain rollback plans. 7.Embed robust cyber- and data-protection controls network segmentation, TLS-encrypted DICOM/HL7 links, and quarterly vulnerability scans, to minimize data-breach risk and ensure regulatory compliance. 8.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 9.Provision of regular updates and reports about the product to DOH upon request. 10.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 11.Compliance with the relevant laws and standards of the Department of Health. 12.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-07-31 | HTA25053 | Iv Watch | Medwing Medical LLC | ivWatch, LLC | United States | Medical Device | Monitoring | 1.Consider the use of the ivWatch Model 400 in healthcare facilities that manage patients receiving peripheral IV therapy, particularly in settings, were early detection of infiltration and extravasation. 2. Define clear patient selection criteria, prioritizing individuals with high-risk infusions, difficult IV access, or a history of infiltration. 3.Develop standard operating procedures (SOPs) covering device setup, sensor placement, monitoring protocols, and clinical response to alerts, and integrate them into routine workflows and EMR documentation. 4.Establish a structured training program for nurses, vascular access teams, and biomedical engineers, including hands-on sessions, alert interpretation, and regular competency refreshers. 5.Integrate ivWatch as an adjunct to routine IV site assessments, ensuring it complements but does not replace clinical evaluation and nursing judgment. 6.Monitor system (Nurse Call System) performance by tracking alert frequency, response times, and outcomes, and use this data to guide quality improvement and reduce alarm fatigue. 7.Emphasize the use of single-use components to minimize infection risk and ensure alignment with infection prevention protocols and manufacturer instructions. 8.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 9.Provision of regular updates and reports about the product to DOH upon request. 10.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 11.Compliance with the relevant laws and standards of the Department of Health. 12.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-07-31 | HTA25058 | MagVenture – R30 | Neurointelligence | Tonica Elektronic A/S | Denmark | Medical Device | Therapeutic | • MagVenture TMS Approved for adoption for the following indications in adult patients suffering from treatment-resistant psychiatric disorders, addiction disorders, or chronic pain conditions: 1. Adults with Depression – MDD/TRD 2. MDD in Adolescents (age 15–21) 3. Adults with OCD 4. Adults with Generalized Anxiety Disorder with criteria for use 5. Adults with PSUD/Addiction – Cocaine with criteria for use 6. Adults with Chronic Pain: to stimulate peripheral nerves for relief of i. chronic intractable ii. post traumatic iii. post-surgical pain. • The therapy should be administered as an adjunct to standard psychiatric treatment, including pharmacological therapy and behavioral therapy. • The system must be operated by trained psychiatrists, neurologists, or certified TMS clinicians familiar with rTMS treatment protocols and safety guidelines. • Pre-treatment screening should be conducted to evaluate seizure risk, presence of cranial implants, or other contraindications prior to initiating therapy. • Treatment parameters (coil type, stimulation intensity, frequency, number of pulses, and session schedule) should follow validated clinical protocols and manufacturer recommendations. • Patient monitoring during treatment sessions is required to ensure safety and detect potential adverse events such as headache, scalp discomfort, or transient neurological symptoms. • The system should be integrated into multidisciplinary care pathways, particularly for treatment-resistant depression, addiction medicine programs, and behavioral health services. • Healthcare facilities should ensure initial training, periodic competency assessment, and adherence to clinical protocols for all staff delivering rTMS therapy. • Clinical outcomes and adverse events should be documented using validated scales (e.g., HAM-D, MADRS, Y-BOCS, addiction severity indices, pain scores) to support clinical governance and real-world evidence collection. Clinical Use Overview: • Population / Intended User: Adult patients suffering from treatment-resistant psychiatric disorders, addiction disorders, or chronic pain conditions. • Indications: 1. Adults with Depression – MDD/TRD 2. Adults with OCD 3. Adults with Chronic Pain:  Peripheral nerves  Post-traumatic,  Post-surgical pain. 4. MDD in Adolescents (age 15–21) 5. Adults with PSUD/Addiction – Cocaine  Adults ≥18 years  Diagnosis of cocaine use disorder  Failure of behavioral and psychosocial interventions  Use as adjunct to structured rehabilitation programs  Subject to outcome tracking and periodic HTA review 6. Adults with Generalized Anxiety Disorder for specific criteria:  Symptoms present for ≥6 months with severe thresholds scores: o Hamilton Anxiety Rating Scale ≥ 18–20 o Generalized Anxiety Disorder 7 ≥ 10–15  Treatment Resistance: Failure of ≥1 or ≥2 antidepressants (often SSRIs or SNRIs) 7. MagVenture TMS Therapy System is indicated as an adjunct for the treatment of Major Depressive Disorder in adolescent patients (age 15–21) • To Be Performed By: Psychiatrists, neurologists, addiction specialists, pain specialists, and trained healthcare professionals. • Clinical Setting: 1. Hospitals 2. Psychiatric treatment centers 3. Addiction treatment clinics 4. Neurology clinics 5. Outpatient mental health centers • Condition of Use: Non-sterile clinical environment; intermittent outpatient therapy delivered across multiple sessions according to clinical treatment protocols. • Contraindications 1. Implanted electronic devices such as pacemakers or neurostimulators located near the head. 2. Metallic implants in the cranial region that could interact with magnetic stimulation. 3. History of seizures or epilepsy, unless medically evaluated and cleared by a specialist. 4. Severe neurological disorders where cortical stimulation may increase risk. 5. Active intracranial pathology or unstable medical conditions. 6. Pregnancy.
2025-07-01 | HTA25049 | Shockwave reducer | Access Medical LLC | Shockwave Medical, Inc. | United States | Medical Device | Therapeutic | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. 4.Shockwave Reducer to be used in patients with debilitating angina and obstructive CAD refractory to optimal medical and revascularization strategies, a reducer device for coronary sinus constriction may be used to improve symptoms, in experienced centers. 5.Continuous collection and reporting of post-market safety and effectiveness data on yearly basis to further monitor real-world effectiveness and safety performance. Population/ Intended User: –In patients with debilitating angina and obstructive CAD refractory to optimal medical and revascularization strategies To be performed by: –Experienced interventional cardiologists Clinical Setting: –Restricted to hospitals with experience in the procedure 6.Restriction to patients with a failed response to optimal medical therapy and not candidates for standard revascularization procedures. 7.Prohibition in patients with a known severe reaction to stainless steel or nickel, chronic renal failure, or severe valvular heart disease, aligning with exclusion criteria noted in clinical evaluations. 8.Periodic Assessments: Requirement that periodic reports be submitted, detailing updates on adverse events, any changes in clinical practice guidelines influencing the device's usage, and the ongoing effectiveness in various patient subgroups to DOH.
2025-07-01 | HTA25028 | OLYMPUS VISERA S VIDEO SYSTEM CENTRE OTV-S500, DEVICE | GULFDRUG LLC | OLYMPUS MEDICAL SYSTEMS CORPORATION , JAPAN | Japan | Medical Device | Diagnostic | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. 4.Deploying the Olympus VISERA S (OTV-S500) is recommended for adoption in medical facilities in Abu Dhabi that focus on minimally invasive and precision-based surgeries. 5.Define clear criteria for patient selection and prioritize procedures that use minimal anesthesia. 6.Develop standard operating procedures (SOPs) for all procedural stages, integrating them into clinical workflows. 7.Develop a structured training curriculum that includes hands-on simulations and periodic refresher sessions for healthcare providers, nursing staff, and biomedical engineers to ensure proper device maintenance and corrective actions. 8.Establish clear inclusion and exclusion criteria, utilize pre-assessment tools, and prioritize awake or minimal-anesthesia procedures for optimal outpatient use. 9.Create comprehensive SOPs for each stage of the procedure, from preparation to post-procedure recovery, and integrate them into the healthcare facility's clinical system. 10.Emphasize the single-use design to minimize cross-contamination and infection risks, aligning with national infection control guidelines. 11.Due diligence on USB usage to be advised. 12.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 13.Compliance with the relevant laws and standards of the Department of Health. 14.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-07-01 | HTA25046 | Dexcom G7-STP-GT-003 | Julphar Drug Store/ DEXCOM INTERNATIONAL LIMITED | DEXCOM | USA | Medical Device | Diagnostics | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. 4.The Dexcom G7 Continuous Glucose Monitoring (CGM) System is a real time, continuous glucose monitoring device 5.indicated for the management of diabetes in persons 2 years and older. 6.Define clear criteria for patient selection for patients with Type 1 and insulin-treated Type 2 diabetes, including children aged 2 years and above and pregnant women, as the Dexcom G7 is approved for use in these populations. 7.Develop a structured training curriculum Including hands-on training for patients on how to apply the sensor and provide guidance to physicians for remote monitoring and patient follow-up. 8.Monitor skin reactions by identifying and managing any local skin irritation or allergic responses. 9.Ensure safe use of Dexcom G7 around medical imaging equipment; do not use the system during MRI, diathermy, or unshielded CT procedures due to potential risks such as device malfunction, inaccurate readings, and loss of alerts by following safety protocols to avoid electromagnetic or thermal interference. 10.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 11.Compliance with the relevant laws and standards of the Department of Health. 12.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-07-01 | HTA25044 | Actim PROM | Sundial Medical Equipment LLC | Actim Oy | Finland | Diagnostic Test | Diagnostics | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. 4.The Actim PROM Intended to be used for prescription use in point of care and clinical labs setting to help diagnose the rupture of fetal membrane (ROM) in pregnant women≥29 weeks gestation who present with signs, symptoms or complaints suggested of ROM. Population/ Intended User: –pregnancy To be performed by: –Healthcare professionals Clinical Setting: –clinical laboratories and point of care Condition of use: –diagnose the rupture of fetal membrane (ROM) in pregnant women≥29 weeks gestation who present with signs, symptoms or complaints suggested of ROM. 5.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 6.Compliance with the relevant laws and standards of the Department of Health. 7.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-07-01 | HTA25042 | Actim Partus | Sundial Medical Equipment LLC | Actim Oy | Romania | Medical Device | Diagnostics | 1.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request. 4.The Actim Partus is intended to be used for the detection of clinical hormones such as IGFBP-1 insulin -like growth factor binding protien-1 in vaginal or cervical samples to aid in the diagnosis of pregnancy complications. Population/ Intended User: –pregnancy To be performed by: –Healthcare professionals Clinical Setting: –clinical laboratories and point of care Condition of use: –diagnosis of pregnancy complications 5.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 6.Compliance with the relevant laws and standards of the Department of Health. 7.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-06-12 | HTA25041 | FebriDx | Babirus Medical Equipment Trading L.L.C. | Lumos Diagnostics | USA | Diagnostic Test | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. To be used in primary and urgent care at the point care settings to aid in differentiating bacterial and viral infections. Further tests like PCR panels or bacterial cultures may be needed for confirmation in complex cases. Approval will be limited to the age 12-64 years old as per the FDA. For use in Pediatric population a proof of concept study in pediatrics is needed before extending the age of approved indication.
2025-06-12 | HTA25037 | Genedrive System; Genedrive CYP2C19 ID Kit; Genedrive CYP2C19 Control Kit | Biocentric | Genedrive Ltd | United Kingdom | Diagnostic Test | Diagnostic | Approval requires compliance with UAE Federal Law No. 49 of 2023 on genetic data storage and transfer. The device must be operated by trained professionals to ensure accurate handling and results. Suitable for emergency and near-patient settings, ensuring quick therapeutic interventions. Required for single-use cartridges; deviations impact assay performance. Contraindications – The test does not detect all CYP2C19 variants, so alternative testing may be necessary for broader genetic profiling.
2025-06-12 | HTA25035 | SENTIERO | HESSA MEDICAL EQUIPMENTS CO LLC | PATH MEDICAL GmbH | Germany | Device | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. Primarily used by audiologists and ENT professionals in clinics, hospitals or mobile settings to detect hearing loss and evaluate auditory function efficiently. Follow manufacturer guidelines for routine calibration and performance checks. Ensure software updates are applied to maintain accuracy and functionality. Keep maintenance and inspection record for traceability and compliance. Sentiero is sufficient for most screenings and diagnostic needs, adding specialized tools can help confirm findings or diagnose more complex conditions.
2025-06-12 | HTA25030 | NËO MONITOR | HESSA MEDICAL EQUIPMENTS CO LLC | eemagine Medical Imaging Solutions GmbH | Germany | System/Software/ AI/App | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. To be used in NICUs to support continuous cerebral monitoring in neonates
2025-06-12 | HTA25032 | AquaBeam Robotic System for Aquablation Therapy | GulfDrug LLC | Procept Biorobitc | USA | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. Approved to be used for the resection and removal of prostate tissue in males suffering from lower urinary tract symptoms (LUTS) due to benign prostatic hyperplasia.
2025-06-12 | HTA25020 | XPERT®COV-2/FLU/RSV PLUS, IVD TEST KIT | Zahrawi Medical Services WLL | Cepheid | USA | Diagnostic Test | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. Using Xpert Xpress SARS-CoV-2/Flu/RSV plus for symptomatic patients. 2. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory.
2025-06-12 | HTA25027 | STANDARD™ Fluorescent Immunoassay (FIA) System (POCT) - F2400 ANALYZER | Global Medical Surgery Equipments and Pharma LLC | SD Biosensor Inc | South Korea | Device | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. To be used for monitoring and diagnosing from the body fluid parameter in clinical settings by healthcare professionals. 2. Used in point of care (POC) and laboratory setting. 3. For infectious diseases, Not a standalone diagnositc definitive test
2025-06-12 | HTA25025 | QScreen (Hearing Screener Device) | HESSA MEDICAL EQUIPMENTS CO LLC | PATH MEDICAL GmbH | Germany | Device | Screening | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. Establish routine inspection and calibration schedules to maintain test accuracy. Implement strict guidelines for usage conditions, particularly in environments with potential electronic interference. Require service and repair to be conducted by authorized partners to preserve device integrity. QSCREEN should be employed alongside other diagnostic tools to provide a comprehensive assessment of hearing health.
2025-06-12 | HTA25036 | GORE® TAG® Thoracic Branch Endoprosthesis (TBE) | ARABIAN MILLENNIUM MEDICAL TRADING LLC | W. L . GORE & ASSOCIATES | USA | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. Restricted to the specified patients categories limited to approved indications. Restricted to centers performing complex cardiological procedures. Applicant to provide DoH updates over the postmarket surviellance yearly along with updates on the running clinical trials: 1- Cont Follow-up IDE Study Subjects 2- GORE TAG TBE Zone 0/1 New Enrollment PAS 3- Continued Follow-up of the IDE Study Report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division in DOH as part of a Real World Data
2025-06-12 | HTA25031 | Genio Bilateral Hypoglossal Nerve Stimulation (HGNS) | Odin Healthcare is the exclusive distributor for Nyxoah in the UAE and represented legally by Pharmalink | Nyxoah SA | Belgium | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. Approved technology with the eligible patients and after exlusion the contraindications mentioned o The Genio® system components are prescription devices. o Use of components not provided by Nyxoah may result in device malfunctioning, damaged devices or increase the risks to the users. o Single-Use Devices o MRI Safety Information should be available The Genio® system is contraindicated for: • Children • Patients with an Apnea-Hypopnea Index (AHI) below 15 or over 65 events/hour • Patients with combined central and mixed AHI > 25% of the total AHI • Pregnant women, or women planning to become pregnant. • Patients with any functional or structural problem that would impair the ability of a hypoglossal nerve stimulator to treat OSA. • Patients with any medical illness or condition that contraindicates a surgical procedure under general anesthesia or that would prevent the implantation of the Implantable Stimulator or the placement of Activation Chip/Disposable Patch • Patients with coagulopathy or requiring anticoagulant medications (such as warfarin, Clopidogrel (Plavix) or similar; prophylactic aspirin not exclusionary) that cannot be safely stopped in the perioperative period. • Patients with major head and neck abnormalities narrowing the airway or the implantation site. • Patients with hypersensitivity to any material of system components • Patients with other Active Implantable Medical Devices (AIMD)
2025-06-12 | HTA25018 | ClearVision Single Use Hysteroscope | Actionmed Medical Equipment Trading LLC | OriGyn Medical Technology (Hangzhou) Co. Ltd | China | Device | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. To be used in outpatient gynecology clinics. 2. Emphasize the single-use design to minimize cross-contamination and infection risks, aligning with national infection control guidelines.
2025-06-12 | HTA25029 | BioBridge Collagen Matrix | Healthpoint | FIBRALIGN CORP | USA | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. To be used in: • Procedures for soft tissue reinforcement, particularly in tendon and ligament repairs. • Procedures for soft tissue repair specifically in lymphatic tissue repair. To be used in hospital setting by Surgeon
2025-06-12 | HTA25026 | Abbott- i-STAT TBI Whole Blood Assay Version 2 - Blood Biomarkers | Abbott Laboratories GMBH | Abbott Point of Care Inc | Canada | Diagnostic Test | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. To be used in emergency and urgent care settings to enhance mTBI (mild traumatic brain injury) evaluation protocols by Integrating the technology into emergency departments and urgent care settings to expedite mTBI assessment and decision-making.
2025-05-12 | HTA25019 | PolyGone Single Use Tissue Resection Device (HR-A-01) | Actionmed Medical Equipment Trading LLC | OriGyn Medical | China | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. The device could be used only by a trained gynecologist only in a well-equiped hospital that is ready to deal with gynecology emergency cases which could be adverse effects by using the device. 2. The hospital must be ready to deal with blood transfusion cases and shock cases. 3. Reporting of the adverse effects of the device 4. As it's a single-use device, providing sterility proof to the hospital is mandatory while supply the device 5. Report to DoH if the device causes adverse effects more than the current standard of care 6. To ensure not using the device in the contraindication cases
2025-05-12 | HTA25022 | Airsmedical Swift MR | KLAS MED LLC | AIRSMED INC | South Korea | System/Software/ AI/App | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. Clinical Adoption: Implement SwiftMR in clinical settings to reduce scan times and enhance image quality. 2. Training: Provide comprehensive training for seamless integration into existing workflows. 3. Regulatory Compliance: Ensure continuous adherence to regulatory standards. 4. Performance Monitoring: Regularly evaluate effectiveness and gather feedback for optimization. 5. R&D: Encourage ongoing research to expand capabilities and improve noise reduction algorithms.
2025-05-12 | HTA25021 | NBCL CONNECT Immunoassay Analyzer | AMICO | NBCL | Netherlands | Device | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. Optimize Patient Selection and Clinical Protocols , Deployment Priorities: 1. Surgical Departments: For intraoperative PTH monitoring during parathyroidectomies. 2. Emergency Departments: For rapid assessment of cardiac biomarkers in acute presentations. 3. Intensive Care Units: Where timely biomarker data can inform critical Conduct Training and Integration into Clinical Workflow - Conduct regular staff training on proper sensor application (Nurses, clinical team) , device calibration(Biomedical Engineers ) , and troubleshooting to maximize accuracy and efficiency.
2025-05-12 | HTA25012 | ILY Telemanipulator Robot | Pioneer Medical Supplies LLC | Sterlab | France | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. Optimize operator selection and use protocols by ensuring proper operator selection for cleanroom or high-risk environments and following validated protocols for robot positioning, task execution, and calibration. 2. Enhance training and integration into workflow by conducting regular training sessions for clinical, technical, and biomedical engineering staff, and integrating robot operations into standard operating procedures (SOPs). 3. As per the Information Security Recommendations, remote access for surgeons should be limited to within the facility network and premises
2025-05-12 | HTA25010 | HbA1c501 Point care of testing | Zahrawi medical services | HemoCue AB | Korea | Diagnostic Test | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1- The device could be used as a tool for diabetes monitoring and screening, provided appropriate quality control. 2- Extra monitoring and screening is required for diagnostic use in special patient population (e.g. pregnancy, hemoglobinopathies) 3- Provide quality assurance systems and ongoing profeciency testing 4- The readings must be aliogned with clinical sysmptoms and does not substitute laboratory testing in case of doubt or clinical disassociation with the readings
2025-05-12 | HTA25008 | Q-Suite (Pantheon and AI Modules) | Qritive Pte Ltd | Qritive Pte Ltd. | Singapore | System/Software/ AI/App | Diagnostic | 1.Q-Suite shall be used as a supportive diagnostic tool in pathology laboratories specializing in prostate cancer to enhance current diagnostic practices. 2.Company to continuously monitor patients’ outcomes to ensure technology’s effectiveness and safety and report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division @Medical Research and Development in DoH as part of a Real World Data. 3.The results from this technology should ONLY be interpreted in correlation with clinical and other laboratory studies; only direct treating physician with expertise in prostate cancer (e.g. urologist, radiation oncologist, medical oncologist etc.) should communicate with a patient the results from the analysis used by this software for prostate cancer. 4.Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 5.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 6.Provision of regular updates and reports about the product to DOH upon request. 7.Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 8.Compliance with the relevant laws and standards of the Department of Health. 9.Compliance with STANDARD ON MEDICAL DEVICE REPORTING (MDR) (DOH/PVE/SD/MDR/V1) of the Department of Health and assign pharmacovigilance focal point for medical product reporting.
2025-04-04 | HTA25017 | Pulmonary Artery Denervation (PADN) | HEARTBEAT MEDICAL SUPPLIES LL | Pulnovo Medical (Wuxi) Co., Ltd | China | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. To be used for treatment of pulmonary arterial hypertension (PAH) (WHO Group 1 pulmonary hypertension) as per the FDA approval for humanitarian use device (HUD) 2.To be perfomed by interventional cardiologists in the cath lab. 3. Company to continuously monitor patients outcomes to ensure technology’s effectiveness and safety and report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division in DOH as part of a Real World Data.
2025-04-04 | HTA25013 | FreeStyle Libre 2 Plus Flash Glucose Monitoring System (Sensor Kit) | Abbott Diabetes Care (ALPHAMED DRUG STORE[Dubai]) | Abbott Diabetes Care | UK | Device | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. As this device is not a new device; we are recommending using the same previous recommendations issued formally to the applicant. 1. The technology may only be used after having been prescribed by a healthcare professional based on the patient's need. 2. Using the technology may only be started after completion of the required training from a qualified instructor. 3. Follow the instructions provided by the manufacturer for the users.
2025-04-04 | HTA25011 | Rbfracture-Radiobotics Aps | Emerging Global Technologies | Radiobotics Aps | Denmark | System/Software/ AI/App | Diagnostic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. It is not intended as a replacement for the clinician’s complete review or their clinical judgment. 2- Assist radiologists by analyzing images and providing results to support the assessment of fractures. 3.Training: Provide comprehensive training for healthcare professionals on its use and limitations. 4. Quality Control: Monitor performance, especially with low-quality images. 5. Guideline Compliance: Adhere to usage instructions and limitations to ensure patient safety.
2025-04-04 | HTA25009 | Vibrant Soundbridge (VIBRATING OSSICULAR PROTHESIS VORP 503 IMPLANT KIT) | CITY PHARMACY COMPANY - SOLE PROPRIETORSHIP L.L.C | MED-ELELEKTROMEDIZINISCHE GERATE GMBH | Austria | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1.VSB should be approved for patients aged 5 years and above with mild to severe SNHL, who are unable to benefit from conventional hearing aids or are unsuitable for cochlear/bone conduction implants. 2.Must involve ENT surgeons, audiologists, and rehabilitation specialists to ensure proper candidate selection, surgery, and post-operative support. 3.Surgical teams should be trained in FMT placement techniques, and follow IFUs and surgical guides to reduce risks and optimize outcomes. 4.Educate patients on maintenance, cleaning, and expected outcomes to maximize satisfaction and device longevity.
2025-04-04 | HTA25007 | SENTEC tCOM+ Transcutaneous Monitor | HESSA MEDICAL EQUIPMENTS CO LLC | UAE AGENT | Swizerland | Device | Monitoring | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. Optimize Patient Selection and Clinical Protocols - Ensure appropriate patient selection, particularly in neonatal, pediatric, and critically adult patients. - Follow guideline-based protocols for sensor site placement, calibration, and monitoring duration to ensure accurate readings and avoid skin irritation or burns. 2. Enhance Training and Integration into Clinical Workflow - Conduct regular staff training on proper sensor application (Nurses, clinical team) , device calibration(Biomedical Engineers ) , and troubleshooting to maximize accuracy and efficiency
2025-04-04 | HTA25006 | AlphaVac Multipurpose Mechanical Aspiration System | AlWasfah Medical Supplies Trading LLC | AngioDynamics Inc | USA | Device | Therapeutic | Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. Provision of regular updates and reports about the product to DOH upon request. 1. Patient Selection & Clinical Guidelines Adherence to use AlphaVac in patients with pulmonary embolism (PE) or deep vein thrombosis (DVT) where systemic thrombolysis is contraindicated or ineffective. 2. Training & Competency Development • Ensure all interventional teams (vascular surgeons, interventional radiologists, cardiologists) receive specialized training in AlphaVac device operation. • Establish standardized protocols for pre-procedural imaging, device setup, and post-procedural monitoring to enhance safety and effectiveness.
2025-04-04 | HTA25005 | Falak Tayyeb Multifunctional Health Monitor | Prome Medical Store | Xiamen Linktop Technology Co.,Ltd | China | Device | Diagnostic | 1. Physicians or healthcare professionals, or personnel who are using the technology should undergo operational or specialized training and required certification/s (whenever applicable). 2.Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3.Provision of regular updates and reports about the product to DOH upon request.
2025-04-04 | HTA24059 | Unified Arrhythmia Diagnostic System-PocketECG III | LIFE MEDICAL EQUIPMENT | Medicalgorithmics S.A. | Poland | Device | Diagnostic | 1. The technology should use by healthcare professionals, including cardiologists, electrophysiologists, and other medical practitioners who specialize in cardiac care. – Population/ Intended User; 18 Years and older. – To be performed by: intended for use by healthcare professionals, including cardiologists, electrophysiologists, and other medical practitioners who specialize in cardiac care. – Clinical Setting: Hospitals., special cardiology center. – Condition of use: Patient who requires monitoring of effect of drugs to control ventricular rate in various atrial arrhythmias. Patients recovering from cardiac surgery are indicated for outpatient arrythmia monitoring. Patients require arrhythmia evaluation of stroke or transient cerebral ischemia, possibly secondary to atrial fibrillation or atrial flutter. Patients who have demonstrated a need for cardiac monitoring e.g. non-life-threatening arrhythmias such as atrial fibrillation atrial flutter, PACs, PSVT) Patients with symptoms that may be due to cardiac arrhythmia such as dizziness, lightheadedness, and dyspnea. Patient with palpitations with or without known arrhythmlas to obtain correlation or rhythm with symptoms
2025-04-04 | HTA24110 | Limusave MT-7500 | Abu Dhabi International Medical Services | Fujifilm | Japan | Device | Diagnostic | 1. The LIMUSAVE MT-7500 should performed by Medical Lab Technology. 2. Continuous monitoring of patient outcomes to ensure technology’s effectiveness and safety. Population/ Intended User: – This device is primarily intended for use in clinical laboratories and healthcare settings were early detection of invasive fungal infections. – To be performed by: – Medical Lab Technology Clinical Setting: – Laboratories Condition of use: – early detection of invasive fungal infections. Exclusion criteria. – Patients who have received certain antifungal treatments or other medications that might affect (1→3)-β-D-glucan levels could be excluded.
2025-04-04 | HTA23050 | SleepUp Digital Therapy | Propharma Medical Supplies | SleepUp Tecnologia em Saúde Ltda | Brazil | System/Software/ AI/App | Therapeutic | 1. SleepUp does not replace physician consultations and is not a diagnostic tool. 2. Medical assessment is required for screening before referral during pregnancy or for individuals with comorbidities. 3.Company to continuously monitor patients outcomes to ensure technology’s effectiveness and safety and report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division in DOH as part of a Real World Data. Population/ Intended User; – Adults with insomnia and sleep-related complaints To be performed by: – Healthcare providers, psychologists, and sleep specialists. Clinical Setting: – Digital/home-based therapy, integrated with telemedicine. Condition of use: – As per manufacturer instructions. Exclusion criteria: – Individuals with severe psychiatric conditions or other medical contraindications.
2025-03-19 | HTA25015 | Toumai® Endoscopic Surgical Robot at Glance ; single port and multiple port. | Cleveland Clinic Abu Dhabi | Shanghai MicroPort MedBot (Group) Co, Ltd. | China | Device | Therapeutic | 1. Approval of The technology Toumai® Endoscopic Surgical Robot Multiport for use in Abu Dhabi , 2. Conditional approval of the Medbot Toumai® Endoscopic Surgical Robot single port robot for the use in clinical research in AD. 3. Healthcare facility to align with the DOH licensing telemedicine requirements. 4. Develop and maintain a comprehensive training program for medical personnel, including surgeons, biomedical and clinical engineers, and the clinical team, to ensure the highest standards of care and safety in operating the endoscopic surgical robot. 5. Ensure the availability of a field trained engineer in Abu Dhabi 24/7 to address any malfunctions, perform preventive and corrective maintenance, and provide on-site technical support in the operating theatre/ rooms as needed. 6. Establish a robust quality monitoring process to track and report any adverse events or unforeseen consequences during usage.
2025-03-17 | HTA24096 | Atellica VTLi Immunoassay Analyzer | One Health LLC | Siemens Healthcare Diagnostics Inc. | Germany | Diagnostic Test | Diagnostic | 1. To be performed as a rapid assessment and risk stratification of suspected AMI patients. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to DOH upon request. Population/ Intended User: – Patients with suspected acute myocardial infarction. – Emergency and outpatient care providers. To be performed by: – Emergency department physicians. – Cardiologists. – Laboratory technicians. Clinical Setting: – Emergency department physicians. – Cardiologists. – Laboratory technicians. Condition of use: – Rapid assessment and risk stratification of suspected AMI patients. – Near-patient testing to expedite clinical decision-making. Exclusion criteria: – Patients with ST-elevation myocardial infarction (STEMI). – Patients requiring extensive laboratory-based cardiac assessments. – Severe hemolysis or lipemia affecting sample integrity. – Patients on anticoagulant therapy that may interfere with hs-cTnI measurement.
2025-03-17 | HTA24079 | Thermalytix 360 | Niramai Middle East medical service LLC formally ASAS medical- Abu Dhabi | Niramai Health Analytix Pvt Ltd | India | System/Software/ AI/App | Diagnostic | 1.It is intended for use as adjunctive diagnostic imaging for thermally significant indications in the breast region, as it only a complementary tool in the clinical examination of the breast, with mammography being the primary screening modality. 2. Company to continuously monitor patients outcomes to ensure technology’s effectiveness and safety and report back to DOH on annual basis in the first 2 years of approval in alignment with Medical Research Division in DOH as part of a Real World Data. 3.. This device is intended for use by qualified healthcare personnel trained in its use. 4. The system is not intended to be used as a thermometry device. 5. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 7. Provision of regular updates and reports about the product to DOH upon request.
2025-03-17 | HTA24078 | CONNEVO KOA, CONNEVO ALI | Burjeel | CONNECTEVE Co. Ltd., 7-8F | South Korea | System/Software/ AI/App | Diagnostic | We recommend approval of using this technology for Market entry with the following conditions: 1- It is not intended as a replacement for the clinician’s complete review or their clinical judgment. 2- Assist radiologists and orthopedic specialists by analyzing knee X-ray images and providing results to support the assessment of osteoarthritis. 2- The final decision should always rest with the interpreting physician, considering the patient’s clinical signs, symptoms, and other relevant information Moreover, DOH has the right to stop the product at any stage if deemed necessary, initial conditions and any subsequent conditions must be satisfied before obtaining final approval. Failure to do so will reflect in provoking the approval.
2025-03-17 | HTA24084 | Veolity | ARJ Medical Supplies LLC | MeVis Medical Solutions, | Germany | System/Software/ AI/App | Diagnostic | 1.It is not intended as a replacement for the clinician’s complete review or their clinical judgment. 2.Radiologists and trained medical imaging professionals.
2025-03-17 | HTA24055 | SIBIONICS GS1 Continuous Glucose Monitoring System | Pharma World L.L.C | Shenzhen SiSensing Co., Ltd., | China | Device | Diagnostic | 1. The technology may only be used after being prescribed by a healthcare professional. 2. Using the technology may only be started after completion of the required training from a qualified instructor. 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Provision of regular updates and reports about the product to DOH upon request.
2025-02-21 | HTA24070 | EYE MAX Single use Cholangioscope incl Harness | Premium Medical Supplies L.L.C | Micro Tech Nanjing Co Ltd, | China | Device | Therapeutic | 1. Establish a quality monitoring process for adverse events and safety concerns. 2. Provide regular updates and reports to the Department of Health upon request. 3. Ensure all training and certification requirements for operators are met.
2025-02-21 | HTA24068 | IVUS Catheter (Intra vascular Ultrasound)/ D01-50 | Heart Beat Medical Supplies LLC | SHANGHAI PULSE MEDICAL TECHNOLOGY, INC. | China | Device | Therapeutic | Establish Clear Criteria for Patient Selection and Usage: o Patient Selection: Define specific criteria for selecting patients who would benefit most from the IVUS catheter, such as those with complex coronary lesions, heavily calcified arteries, left main disease, and in-stent thromboses. o Exclusion Criteria: Clearly outline exclusion criteria to ensure patient safety, including conditions like bacteremia, sepsis, major coagulation system abnormalities, severe hemodynamic instability, and total occlusion. o Usage Protocols: Develop detailed protocols for the use of the IVUS catheter during coronary interventions, ensuring that it is used appropriately and effectively. 2. Implement Comprehensive Training and Certification Programs: o Training Programs: Create and maintain extensive training programs for medical personnel, including interventional cardiologists, radiologists, and catheterization lab technicians. Ensure they are fully trained in the use of the IVUS catheter and the interpretation of its imaging results. o Certification: Implement certification 1. processes to validate the competency of healthcare professionals in using the IVUS catheter, ensuring high standards of care and safety. o Ongoing Education: Provide continuous education and training updates to keep medical personnel informed about the latest advancements and best practices in intravascular ultrasound technology. 3. Develop Robust Monitoring and Follow-Up Protocols: o Patient Monitoring: Establish protocols for regular monitoring of patients during and after the use of the IVUS catheter to assess the response to treatment and manage any potential side effects. o Long-Term Follow-Up: Develop a long-term follow-up plan to track patient outcomes, gather additional data on the effectiveness of the IVUS catheter, and identify any long-term complications or benefits. o Data Collection: Implement systems for collecting and analysing data on the use of the IVUS catheter, contributing to continuous improvement in patient care and technology utilization.
2025-02-21 | HTA24090 | HeRO | MDC CARE MEDICAL TRAD ING LLC. | Medical Predictive Science Corporation, | USA | System/Software/ AI/App | Diagnostic | 1. The technology should be performed by licensed health care practitioner in neonatal or paediatric ICU. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to
2025-02-21 | HTA24085 | MeVis LiverSuite | ARJ Medical Supplies LLC | MeVis Medical Solutions, | Germany | System/Software/ AI/App | Diagnostic | 1. Radiologists and trained medical imaging professionals. 2. Patient Health Information to be anonymized as per DOH Ino Sec 3. Right to audit by regulatory body (DOH).
2025-02-21 | HTA24089 | Navina Classic System | Alphamed General Trading LLC | Wellspect Healthcare, | Sweden | Device | Therapeutic | 1. Provide user and patient training to ensure optimal use. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to DOH upon request.
2025-02-21 | HTA24066 | EnSite X EP System | Abu Dhabi International Medical Services | Abbott Laboratories | USA | Device | Diagnostic | 1. Establish a quality monitoring process for adverse events and safety concerns. 2. Provide regular updates and reports to the Department of Health upon request. 3. Ensure all training and certification requirements for operators are met.
2024-11-20 | HTA24080 | VARIPULSE BI-DIRECTIONAL CATHETER/D141201 | Johnson & Johnson Middle East INC-DHCA Dubai | BIOSENSE WEBSTER INC. | USA | Device | Therapeutic | 1. VARIPULSE Bi-Directional Catheter/D141201 should be used in patients with drug-refractory paroxysmal atrial fibrillation (PAF) 2. The procedure should be performed only in specialized centers with electrophysiologists experienced in ues. 3. Physicians using the VARIPULSE Bi-Directional Catheter should undergo specific training and certification in the use of PFA technology. To be used in - Primary therapy: As a first-line ablation option, especially in patients who cannot tolerate antiarrhythmic medications or as an alternative to traditional thermal ablation techniques (RF or cryoablation). - Secondary therapy: In cases requiring repeat ablation after failed thermal approaches, leveraging its tissue-selective pulsed-field ablation (PFA) mechanism.
2024-10-30 | HTA24056 | BIOFLEX MultiPort System - photobiomodulation (PBM). | A G B L LABORATORY AND MEDICAL EQUIPMENT TRADING L.L.C | Meditech International Inc | Canada | Device | Therapeutic | 1. Technology should apply in healthcare professionals’ hospital. 2. The practitioner must be a Certified Bioflex Laser Specialist. To be used in: Population/ Intended User: Adults (18+) To be performed by: BIOFLEX® MultiPort System is used by trained healthcare professionals Clinical Setting: Hospitals, clinics or ambulatory centers. Condition of use: - relief of neck and shoulder pain. - relief of muscle and joint pain, arthritis, muscle spasm, relieving stiffness and promoting relaxation of muscle tissue. Exclusion criteria: - patients with lupus and other diseases which may be stimulated by light. - history of keloids, recent excessive sun exposure and/or artificial tanning at the treatment area. - recent patients using photosensitive medications or herbs. - Pregnant.
2024-10-30 | HTA24062 | Coloplast TFL Drive - Thulium Fiber Laser Device & Fibers | Pharmatrade LLC | Coloplast Corp | Italy | Device | Therapeutic | 1. Ensure that healthcare professionals, including surgeons and operating room staff, undergo specialized training to handle and operate the Thulium Fiber Laser device To be used in The Coloplast TFL Drive is used in urological surgeries, providing improved outcomes for urinary stone management, BPH treatments, and tumour removal. - Patients with Urinary tract system stone accumulation. - Men’s 50 years age and above with BPH. - Patients with tumours in the soft tissue e.g Bladder.
2024-10-21 | HTA24041 | Recor Paradise ultrasound | Basic Medical Equipments LLC | ReCor Medical | USA | Device | Therapeutic | 1. The procedure is performed in the Cath Lab under fluoroscopy guidance via the right femoral artery under deep sedation and may be performed as a day case procedure in experienced Hospital /centre. 2.Condition of use: -Renal denervation may be of use in adults with uncontrolled or resistant hypertension (triple drug therapy, one of which is a diuretic) and may be an option for patients unable to tolerate long-term medications at the necessary doses or who cannot tolerate medications at all. -There may be a role for renal denervation in patients who simply do not wish to be on medication. Patients must be fully informed about the risks and potential benefits of renal denervation when compared to drug therapy. -Renal denervation should be offered mainly to patients with high global cardiovascular risk, for whom intensive BP treatment may have a particularly pronounced benefit. Hospital offering renal denervation should include a multidisciplinary team of hypertension specialists and expert interventionalists who are trained specifically for renal denervation and whose outcomes are closely monitored. –Exclusion criteria: – Renal arteries diameter<3mm and >8mm. – Renal artery fibromuscular disease (FMD). – Stented renal artery. – Renal artery aneurysm. – Renal artery diameter stenosis >30%. – Pregnancy. – Presence of abnormal kidney (or secreting adrenal). Tumors.
2024-10-21 | HTA24060 | FARAPULSE (PFA) system Pulsed Field Ablation Generator | Al Khayyat Investments (AKI) | Boston Scientific | USA | Device | Therapeutic | 1. FARAPULSE™ PFA system should be used in patients with paroxysmal atrial fibrillation (PAF). 2. The procedure should be performed only in specialized centers with electrophysiologists experienced in catheter ablation techniques. 3. Physicians using the FARAPULSE™ system should undergo specific training and certification in the use of PFA technology. To be used in : ▪Population/ Intended User: –Patients with paroxysmal atrial fibrillation who may not have responded well to antiarrhythmic drugs. –The FARAWAVE Catheter is indicated for the isolation of pulmonary veins in the treatment of drug-refractory, recurrent, symptomatic Paroxysmal Atrial Fibrillation (PAF). ▪To be performed by: –Cardiologists specializing in electrophysiology. ▪Clinical Setting: –Cardiovascular EP procedural settings. ▪Condition of use: –For patients with AF where antiarrhythmic drugs or other forms of ablation are not effective or appropriate. ▪Exclusion criteria: –Any conflicting medical issues or comorbidities that increase risk or contraindicate the use of PFA.
2024-09-26 | HTA24042 | Food intolerance, IgG kit | PureLab | MacroArray Diagnostics | Austria | Product/Kit | Diagnostic | 1. To be done for patients who have food allergy , or food intolerance. 2. Patient who have allergy symptoms but not confirmed yet by lab test
2024-09-26 | HTA24043 | Digital Theraputics - Edupression | Takalam L.T.D | SOFY GmbH | Austria | System/Software/ AI/App | Therapeutic | 1. Approval for the English app. If an Arabic version is to be released, this needs a validation study to be submitted before the release. 2. The technology should be performed by psychotherapists. Provision of regular updates and reports about the product to DOH upon request. To be used in: ▪Population/ Intended User: Age from 18- 60 ▪To be performed by Doctor or psychotherapist ▪Clinical Setting: Hospital, Psychotherapy centers ▪Condition of use: use of digital therapy programs as monotherapy for mild depression and as adjunct therapy for moderate to severe depression. ▪Exclusion criteria: –The use of edupression.com® is therefore not indicated in the presence of suicidal thoughts or bipolar disorder or psychotic symptoms in the context of schizophrenia, a severe depressive episode with psychotic symptoms, a schizoaffective disorder, delusional disorder, or other disorders with psychotic symptoms. –patients who are unable to follow the online psychoeducation offerings for various reasons (e.g. intellectual disability, alcohol or drug intoxication, visual impairment, etc.). –edupression.com® is also not suitable for patients who cannot use a digital device such as a computer/tablet/smartphone or who do not have the necessary language skills, which may apply to some geriatric patient groups or patients with a Native language that does not exist in the language range of edupression.com® or has a lack of foreign language skills applies.
2024-09-26 | HTA24050 | Cooled Radiofrequency Ablation | Health Point Hospital | Avanos | USA | Device | Therapeutic | 1. Implementation of a structured training program for healthcare providers on the use of COOLIEF* technology and must be certified in its operation and safety protocols. 2. Ensure that all safety precautions and guidelines provided by Avanos are strictly followed to prevent risks such as RF burns and unintended stimulation. 3. Continuous monitoring of patient outcomes to ensure the technology’s effectiveness and safety. Population/ Intended User: Patients with chronic pain conditions, particularly those unsuitable for surgery. ▪To be performed by: Trained and certified healthcare professionals specializing in pain management. ▪Clinical Setting: Outpatient clinics, pain management centers. ▪Condition of use: To be used in accordance with manufacturer guidelines and clinical best practices.
2024-08-15 | HTA24028 | Ionic Radiofrequency Generator | Al Amin Medical Instruments LLC | Abbott | USA | Device | Therapeutic | 1. Ensure that all healthcare professionals using the IonicRF™ Generator are adequately trained and certified in its operation and safety protocols. 2. Ensure that all safety precautions and guidelines provided by Abbott are strictly followed to prevent risks such as RF burns and unintended stimulations. This makes it a critical option for patients looking for effective pain relief without the need for more invasive surgical procedures. Management of various chronic pain, including: • Neck • Upper and lower back • Sacroiliac joint • Facial pain
2024-08-15 | HTA24030 | Cardio Explorer | lfaDiag | Exploris Health AG | Switzerland | System/Software/ AI/App | Diagnostic | 1. An aid on diagnostic in the field of Non-invasive CAD evaluation minimizing the need for invasive tests by providing clearer non-invasive diagnostic outcomes. 2. Tailored diagnostics supporting a personalized approach to CAD diagnosis and treatment. 3.The software will not replace the clinical judgment and decision of the treating physician. 4.The sofware will not replace the clinical judjment and decion of the treating physician. 5.Prioritize transparency in AI decisions. 6. Ensure data privacy and integrity. 7. Require training for users for accurate data input and interpretation. 8. Monitor false positives and negatives. 9. Implement strict access controls. 10. Set up a feedback mechanism for continuous improvement. 11. Manufacturer to validate algorithm accuracy periodically.
2024-08-15 | HTA24023 | LINX™ REFLUX MANAGEMENT SYSTEM | Johnson and Johnson Middle East Inc. | TORAX MEDICAL INC,UNITED STATES OF AMERICA | USA | Device | Therapeutic | LINX is primarily indicated for patients diagnosed with Gastroesophageal Reflux Disease (GERD) who continue to have chronic symptoms despite maximum medical therapy, including those intolerant to or seeking an alternative to proton pump inhibitors (PPIs). 1. Surgeons performing the LINX® procedure should be specifically trained in laparoscopic anti-reflux surgery and familiar with the unique aspects of the LINX® device to ensure proper placement and minimize the risk of complications. 2. Patients with specific contraindications, such as large hiatal hernias (greater than 3 cm), Barrett’s esophagus, or severe esophagitis (Grade C or D), should be excluded from receiving the LINX® Reflux Management System. 3. Ensure that patients with a history of allergic reactions to titanium, stainless steel, nickel, or ferrous materials are thoroughly screened before considering the LINX® device, as the sizing tool used during the procedure contains nickel. Patients with allergies to these materials should be excluded from using the LINX® system 4. Ensure robust monitoring for adverse events by establishing a regular follow-ups for patients with the LINX® device to promptly identify and manage any complications, such as device migration or dysphagia, ensuring patient safety and timely intervention. 5. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 6. Provision of regular updates and reports about the product to DOH upon request.
2024-07-17 | HTA24018 | Elecsys IL-6 (Interleukin‑6) 09015604190 - 09015612190 | PHARMATRADE LLC | Roche Diagnostics | Germany | Diagnostic Test | Diagnostic Test | The technology to be performed by Medical lab Technologist.
2024-07-11 | HTA24035 | Heavy Ion Therapy | Cleveland Clinic Abu Dhabi | Toshiba Energy Systems & Solutions Corporation | Japan | Device | Therapeutic | 1.Establish clear criteria for selecting patients, focusing on those with tumors resistant to conventional therapies or located near critical organs, and define exclusion criteria to ensure the safety and suitability of patients for Heavy Ion Therapy. 2.Implement protocols for regular patient monitoring during and after treatment to assess response and manage side effects and develop a long-term follow-up plan to track patient outcomes and gather additional data on the effectiveness of the therapy. 3.Create and maintain a comprehensive training program for medical personnel, including radiation oncologists, biomedical and clinical engineers, medical physicists, and dosimetrists, to ensure the highest standards of care and safety in administering Heavy Ion Therapy.
2024-07-01 | HTA23073 | Interfyl® Human Connective Tissue Matrix | KIZAD | Celularity | USA | Product/Kit | Therapeutic | Population/ Intended User; – Suited for a variety of surgical and wound applications where there is a need to replace or supplement damaged or inadequate integumental tissue. – Ideal solution for tunneling and satellite wounds To be performed by: – Surgeons. Clinical Setting: – Hospitals, special surgery centers. Condition of use: – augmentation of deficient/inadequate soft tissue and treatment of deep dermal wounds; Tears. – surgical wounds; soft tissue voids as a result of tunneling wounds, fistula tracts. – dermal undermining including those with exposed vital structures (bone, tendon, ligament, or nerve). – Diagnosis with end stage joint arthritis. – exhausted all forms of conservative treatment. – elected to undergo foot and/or ankle arthrodesis. Exclusion criteria: – use of post operative bone stimulator . – history of active target joint infection in 6 months prior to surgery. – malignant neoplasm. – inadequate blood perfusion
2024-07-01 | HTA23070 | Human Amniotic Membrance Allograft: 1. BIOVANCE , 2. 3L BIOVANCE , 3. BIOVANCE 3L Ocular | KIZAD | Celularity | USA | Product/Kit | Therapeutic | Select patients based on clinical need, prioritizing those with chronic non-healing wounds and do o not use in patients with active infections, or inadequate blood perfusion. Population/ Intended User; – Suited for a variety of surgical and wound applications where there is a need to replace or supplement damaged or inadequate integumental tissue. – Ideal solution for tunneling and satellite wounds To be performed by: – Surgeons. Clinical Setting: – Hospitals, special surgery centers. Condition of use: – augmentation of deficient/inadequate soft tissue and treatment of deep dermal wounds; Tears. – surgical wounds; soft tissue voids as a result of tunneling wounds, fistula tracts. – dermal undermining including those with exposed vital structures (bone, tendon, ligament, or nerve). – Diagnosis with end stage joint arthritis. – exhausted all forms of conservative treatment. – elected to undergo foot and/or ankle arthrodesis. Exclusion criteria: – use of post operative bone stimulator . – history of active target joint infection in 6 months prior to surgery. – malignant neoplasm. – inadequate blood perfusion
2024-07-01 | HTA24014 | Multiplex Real time PCR Kit for Twenty-three Respiratory Pathogens | Instamed Medical Equipment Trading LLC | UniMedica | China | Diagnostic Test | Diagnostic | 1.Perform the test in local laboratories equipped with specialized equipment and expertise. 2.Ensure that Standard Operating Procedures (SOPs) are maintained for sample collection, processing, and analysis. 3.Provide comprehensive training for laboratory personnel. Population/ Intended User; – Patients require rapid and accurate identification of respiratory pathogens. Exclusion criteria: – Recent infections, especially those targeted by the PCR test, might impact results due to residual nucleic acids. Patients should provide a history of recent illnesses. – Conditions that affect the immune system (e.g., HIV/AIDS, autoimmune disorders) or chronic infections might influence test results. Patients should disclose their complete medical history, including chronic illnesses. – Immunocompromised patients might have altered pathogen loads or atypical presentations of infections, which could influence PCR results.
2024-06-04 | HTA2416 | M42 Biogenix Pharmacogenomics report powered by PharmCAT | M42 | M42 | UAE; powered by PharmCAT | System | Therapeutic | Population/ Intended User; – The Report provides analysis of 23 patient’s genes and determines individual rate of metabolism for each of 128 drugs for the patient. To be performed by: – Health care providers – Clinical Setting: – Hospitals, clinical settings. Condition of use: – patients using long-term medications. – individuals with chronic ailments. – patients prone to side effects or adverse reactions.
2024-05-24 | HTA24021 | Advanced Hybrid Closed Loop MiniMed 780G Insulin Pump System | MEDTRONIC – META | MEDTRONIC | USA | Device | Therapeutic | 1. Advanced Hybrid Closed Loop MiniMed 780G Insulin Pump System,Version: MiniMed 780G. 2. Recommended for patients with type 1 adults and children with poor glycemic control and unable to reach HbA1c and time in Range targeted levels. 3. Using the technology may only be started after completion of the required training from a qualified instructor. 4. Follow the instructions provided by the manufacturer for the users. 5. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees and patients. 6. Provision of regular updates and reports outcome about the product to DOH upon request. Population, setting and intended user for Technology “Advanced Hybrid Closed Loop MiniMed 780G Insulin Pump System” ▪ Population/ Intended User: – Type 1 adults and children with poor glycemic control and unable to reach HbA1c and Time in Range targeted levels. – Patients with severe glycaemic variability; Patients with frequent episodes of hypoglycemia. – Recurrent diabetic ketoacidosis (DKA)/ recurrent hospitalizations. – Dawn Phenomenon; Gastroparesis. – Patient preference, meal timing flexibility and normalization of lifestyle. – Low insulin Requirements (not easily measured by via a syringe) Inability to self-administer insulin. ▪ To be performed by: – Healthcare professionals – Trained individual with diabetes & Trained parents whose children have diabetes ▪ Clinical Setting: – Hospitals & Clean environment ▪ Condition of use: – As per manufacturer instructions. ▪ Exclusion criteria: – People who are not able to be in regular contact with their healthcare professional. – People with skin that does not tolerate adhesive pads. – People who are not able to notice alarms because of physical limitations.
2024-05-21 | HTA23067 | Chocolate Touch Paclitaxel-Coated PTA Balloon Catheter | G Vascular USA, LLC | TriReme Medical, LLC | USA | Device | Therapeutic | 1.Approval on the use of Chocolate Touch Paclitaxel-Coated PTA Balloon Catheter Model TUAA-BBB-XXYYY OTW (AA- guidewire, BB- catheter length, XX balloon diameter, YYY – balloon length) 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to DOH upon request. 4. Any other documents or information requested regarding the product and cost to finalize the approval process. Population, setting and intended user for Technology “Chocolate Touch PaclitaxelCoated PTA Balloon Catheter”. ▪ Population/ Intended User: The Chocolate Touch DCB is used when arteries in the thigh and knee are narrowed or blocked because of peripheral artery disease/ is indicated for percutaneous transluminal angioplasty of de novo or restenosis lesions in length in native femoral or popliteal arteries. ▪ To be performed by: By trained healthcare professionals ONLY in interventional cardiology and vascular surgery operation theatres ▪ Clinical Setting: Hospitals, special surgery canters. ▪ Condition of use: As per Manufacturer directions. ▪ Exclusion criteria:The Chocolate Touch DCB should not be used in: • Arteries that carry blood to the heart (coronary), kidneys (renal), brain (cerebrovascular), or branch off from the largest artery (supra-aortic). • Patients with a known hypersensitivity (allergy) to paclitaxel or drugs with similar characteristics as paclitaxel. • Patients who cannot take recommended medicines that thin the blood and prevent blood clots. • Women who are breastfeeding, pregnant, or intend to become pregnant; or men intending to father children. • If a doctor decides that a patient’s narrowing, or blockage will prevent proper placement of the delivery system
2024-05-21 | HTA23072 | Endotest | Aurora Drug Store L.L.C | DNA GENOTEK INC | Canada | Diagnostic Test | Diagnostic | 1. Approval on Endoset version OME-505.005+ACP119 for the use on symptomatic women between the ages of 18-43. 2. Information Security conditionally approved as following. • Patient Health Information to be anonymised. • VIP and Royal data to be excluded. • The requestor to identify solutions within UAE and compliant with applicable Legal and Regulatory requirements. • Right to audit by regulatory body (DOH) 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Provision of regular updates and reports about the product to DOH upon request. ▪ Population/ Intended User. – A non-invasive diagnostic tool that allows for the early detection of endometriosis ▪ To be performed by: – By Healthcare Professional only ▪ Clinical Setting: – Hospitals, clinic, etc. ▪ Condition of use: – As per Manufacturer instruction ▪ Exclusion criteria: – Other conflicting medical issues
2024-05-10 | HTA24004 | DeepCatch | MEDICAL IP Co., Ltd. | MEDICAL IP Co., Ltd. | Republic of Korea | Device | Diagnostic | Population, setting and intended user for Technology “DeepCatch- Version.1.2.0.me” ▪ Population/ Intended User; – Patients with chronic conditions or undergoing cancer treatment. – Individuals at risk for diseases like osteoporosis or muscle atrophy. ▪ To be performed by: – By Radiologist ▪ Clinical Setting: – Utilize this technology in oncology centers and imaging departments ▪ Condition of use: – Successful Clinical Validation of Prognostic Predictions for Specific Conditions (Muscle Atrophy, Osteoporosis, Hepatitis B, Metabolic Syndrome, obesity, etc.) – Lung cancer – Cervical Cancer. – Ovarian cancer. – Sarcopenia. – Diabitise. ▪ Exclusion criteria: – Pregnant women
2024-04-11 | HTA24010 | Edison System/Histotripsy | Al Naghi Medical Co. LLC | HistoSonics, Inc. | USA | Device | Therapeutic | 1. Approval of Edison System/Histotripsy, SW Version 2.1.X intended to be used for liver tumor. 2. It should be performed by skilled and experienced medical professionals that completed a comprehensive training program. 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Provision of regular updates and reports about the product to DOH upon request. Population, setting and intended user for Technology “Edison System/Histotripsy” ▪ Population/ Intended User; – Individuals with Liver Tumors. – Liver patients seeking non-invasive treatment options. – Individuals with Tumors in Locations Difficult to Treat Surgically ▪ To be performed by: – Skilled and experienced medical professionals that completed a comprehensive training program. ▪ Clinical Setting: – Hospitals specialized in cancer treatment for the indication. ▪ Condition of use: – As per the manufacturer’s instructions. ▪ Exclusion criteria: – Not suitable for patients who recently received live vaccines.
2024-03-29 | HTA-24005 | High-Intensity Focused Ultrasound-(HIFU) -SONOVEIN HD THC 903359-B | Healthpoint Hospital | Theraclion | france | Device | Diagnostic | Population, setting and intended user for Technology “High-Intensity Focused Ultrasound (HIFU) Version - SONOVEIN HD THC903359-B” Population/ Intended User; – varicose veins – incompetent perforating veins – active venous ulcers – venous impotence To be performed by: – Specialist traind in duplex ultrasonogrphy. Clinical Setting: – Hospitals, Specialized clinics. Condition of use: – All patient affected by ALL symptomatic grades of venous insufficiency. Exclusion criteria: – Pregnancy . – Infection at the Treatment Site.
2024-03-13 | HTA23024 | CardioMems CM 2000 | Cleveland Clinic Abudhabi | Abbott | USA | Device | Therapeutic | 1. Approval on using the CardioMems CM2000. 2. Utilizing the device is limited to patients who fall under the specific criteria in medical centers with heart failures services; o Class ll or class lll heart failure. o Patients with past heart failure hospitalizations or have elevated natriuretic peptides (or both). 3. Required a proper training and education for both Patients and healthcare providers. 4. Conditional Approval by Abu Dhabi Healthcare Information Security Program until 21/02/2025. 5. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 6. Provision of regular updates and reports about the product to DOH upon request. 7. Any other documents or information requested regarding the product and cost to finalize the approval process. Population, setting and intended user for Technology “CARDIOMEMS CM2000 Wireless Pulmonary Hemodynamic Monitoring” ▪ Population/ Intended User; CardioMEMS is particularly beneficial to patients with: – Class ll or class lll heart failure. – Past heart failure hospitalizations. – Fluid volumes that are difficult to assess or manage. – Advanced limitation of physical activity due to fatigue and shortness of breath. ▪ To be performed by: – By interventional cardiologist. ▪ Clinical Setting: – Medical centers with Heart Failure services. ▪ Condition of use: – As per the specific criteria. ▪ Exclusion criteria: – The CardioMEMS HF System should not be used in people who are unable to take bloodthinning (anticoagulant) or anti-clotting (antiplatelet) medications for one month after receiving the implant. – The following patients may not be appropriate for implantation of the CardioMEMS™ HF System2: • Patients with an active infection. • Patients with a history of recurrent (> 1) pulmonary embolism or deep vein thrombosis. • Patients unable to tolerate right heart catheterization. • Patients with a Glomerular Filtration Rate (GFR) < 25 ml/min who are non-responsive to diuretic therapy or who are on chronic renal dialysis. • Patients with congenital heart disease or mechanical right heart valve(s). • Patients with known coagulation disorders. • Patients with a hypersensitivity or allergy to aspirin, and/or clopidogrel. • Patients who have undergone implantation of a Cardiac Resynchronization Device (CRT) within the past three months. • If the patient’s BMI is greater than 35, measure the patient’s chest circumference at the axillary level. If the chest circumference is > 165 cm, sensor implantation should not occur
2024-03-13 | HTA23063 | Symplicity Spyral™ Multi-Electrode Renal Denervation (RDN) System | Medtronic | Medtronic | Republic of Ireland | Device | Therapeutic | Population, setting and intended user for Technology “Symplicity Spyral™ MultiElectrode Renal Denervation (RDN) System.” ▪ Population/ Intended User; – This treatment is performed in patients suffering from resistant/uncontrolled hypertension. ▪ To be performed by: – By interventional cardiologist: ▪ Clinical Setting: – Hospitals, special surgery centers ▪ Condition of use: – As per Physician instructions ▪ Exclusion criteria: RDN should be avoided in patients with an unsuitable anatomy of the renal artery/access site or any condition that would increase the risk of the procedure, such as a bleeding disorder
2024-03-06 | HTA23076 | Oxford Immunotec T-SPOT | Al Fahim Healthcare Solutions L.L.C-Branch of Abu Dhabi | Oxford Immunotec | UK | Device | Diagnostic | 1. Approval of Oxford Immunotec T-SPOT. TB test, version: TB.300. 2. The T-SPOT.TB test should be performed by laboratory technicians who are skilled in handling blood samples and conducting immunological assays in a clinical or laboratory setting. 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Provision of regular updates and reports about the product to DOH upon request. Population, setting and intended user for Technology “Oxford Immunotec T-SPOT. TB test” ▪ Population/ Intended User; – Individuals with possible exposure to TB. – Healthcare workers exposed to TB patients. – Immigrants from TB-endemic regions. – Immunocompromised Individuals. – Patients Before Initiating Immunosuppressive Therapy. – Residents and employees of high-risk settings such as prisons, homeless shelters, or nursing homes. ▪ To be performed by: – Laboratory technicians who are skilled in handling blood samples and conducting immunological assays. ▪ Clinical Setting: – Hospitals, clinics, or specialized laboratories are equipped to handle blood samples and conduct immunological assays. ▪ Condition of use: – As per the manufacturer’s instructions. ▪ Exclusion criteria: – Not suitable for patients who recently received live vaccines.
2024-03-03 | HTA23074 | Qxr - CHEST X-Ray -V3.2 | Pulse townhouses company | Qure.ai Technologies Pvt Ltd. | India | Device | Diagnostic | Population, setting and intended user for Technology “Qxr” Population/ Intended User; qXR can identify upto 29 abnormalities : – Atelectasis, Blunted Costophrenic angle, Calcification, Cardiomegaly, Cavity, Consolidation, Degenerative Spine Changes, Hyperinflation, Fibrosis, Opacities, Pleural Effusion, Hilar Prominence, Reticulo Nodular Pattern, Scoliosis, Tracheal Shift, Tuberculosis, Nodules, Pneumothorax, Covid-19 risk, Elevated Hemidiaphragm, Pneumoperitoneum, Rib Fractures, Mediastinal Widening, Linear Opacities, Presence of Tracheal Tube, Presence of Gastric Tube, Placement of Tracheal Tube, Placement of Gastric Tube, Lung Nodule Malignancy To be performed by: – for use by licensed medical practitioner/ radiologist. Clinical Setting: – Hospitals Condition of use: – 6 years and above who have a chest X-ray as part of their clinical workup Exclusion criteria: – Lateral X-rays – X-rays that do not contain the entire lung field. – X-rays of other body parts. – X-rays of patients below 6 years of ag
2024-02-29 | HTA23047 | Microwave Thyroid Ablation | Mediclinic Hospital(Al Noor) | ECO | China | Device | Therapeutic | Population, setting and intended user for Technology “Microwave Thyroid Ablation” ▪ Population/ Intended User; – Patients with benign Thyroid Nodules ▪ To be performed by: – By interventional radiology team. ▪ Clinical Setting: – Hospitals, special surgery centers. ▪ Condition of use: – As per manufacturer guidance ▪ Exclusion criteria: – Other conflicting medical issues
2024-02-26 | HTA23080 | Endofresh Endoscopy | Hayat Biotech LTD. | DAICHUAN MEDICAL | China | Device | Diagnostic | 1. Approval for Endofresh Disposable Digestive Endoscope System, Device version/model number XZING-W200B (Disposable GI Colonoscope)/ XZING-W200B (Disposable GI Colonoscope) / XZING-S2 (Camera System). 2. Using the technology may only be started after completion of the required training from a qualified instructor. 3. Should only be performed in highly selected patients on a case-by-case basis such as those who are immunocompromised or have MDROs. 4. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 5. Provision of regular updates and reports about the product to DOH upon request. Population, setting and intended user for Technology “Endofresh Disposable Digestive Endoscope System, Device” ▪ Population/ Intended User; – Recommended for high-risk patients where infection control is paramount. ▪ To be performed by: – Trained medical professionals. ▪ Clinical Setting: – Hospitals, outpatient clinics, and any healthcare settings equipped for endoscopic procedures. ▪ Condition of use: – The system is designed for single-use applications, thereby ensuring a sterile device for each patient. It's particularly useful in settings with high patient turnover or limited sterilization facilities. Exclusion criteria: – Not suitable for complex endoscopic procedures requiring advanced endoscopic tools or capabilities that reusable endoscopes offer.
2024-01-29 | HTA23081 | FreeStyle Libre 2 | Abbott Laboratories GmbH – Diabetes Care Division | ABBOTT - DIABETES CARE | Ireland | Device | Diagnostic | -Recommended for patients of pre-diabetic and uncontrolled diabetic patients with an A1c >7% regardless of their medication. -The technology may only be used after being prescribed by a healthcare professional. -Using the technology may only be started after completion of the required training from a qualified instructor. -Follow the instructions provided by the manufacturer for the users. -Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees and patients. -Provision of regular updates and reports outcome about the product to DOH upon request. Population, setting and intended user for Technology “FreeStyle Libre 2” ▪ Population/ Intended User; – Individuals with diabetes who require regular monitoring and management of their blood glucose levels. – Recommended for patients of pre-diabetic and uncontrolled diabetic patients with an A1c >7% regardless of their medication ▪ To be performed by: – Healthcare professionals – Trained individual with diabetes – Trained parents whose children have diabetes ▪ Clinical Setting: – Hospitals – Clean environment ▪ Condition of use: – As per manufacturer instructions. ▪ Exclusion criteria: – People who are not able to be in regular contact with their healthcare professional. – People with skin that does not tolerate adhesive pads. – People who are not able to notice alarms because of physical limitations.
2024-01-24 | HTA23058 | NeuRx DPS® Diaphragm Pacing System | Reyada Home Health Care Services L.L.C | Synapse Biomedcial Inc. | USA | Device | Therapeutic | 1. Approval on the use of NeuRx DPS System; version 20-0030 limited to the following criteria: - Only for patients +18 with stable, high spinal cord injuries with stimulable diaphragms. - In specialized healthcare facilities with acute care units for spinal cord injuries and specialized rehabilitation centres with traumatic spinal Cord Injury (SCI) well defined care program. - In the presence of well-trained nurse and skilled rehabilitation team during the activation of the device. 2. For cost effectiveness and coverage of the technology, the request needs to be submitted by the provider for reimbursement through empanelment process. 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Provision of regular updates and reports about the product to DOH upon request. Population, setting and intended user for Technology “NeuRx DPS® Diaphragm Pacing System” ▪ Population/ Intended User; – This device is intended for use in patients with stable, high spinal cord injuries with stimulable diaphragms, but who lack control of their diaphragms. – The NeuRx Diaphragm Pacing System is intended for use in patients 18 years and older with stable high spinal cord injuries who lack control of the diaphragm but who will benefit from diaphragm stimulation to help them breathe without relying on a mechanical ventilator. The system can support patient breathing continuously for 4 or more hours each day. ▪ To be performed by: By skilled surgeon, monitored by well-trained nurse and in the care of skilled rehabilitation team. ▪ Clinical Setting: Hospitals, special surgery centers. ▪ Condition of use: – The NeuRx DPS should not be used for patients with diaphragms that cannot be stimulated. – The NeuRx® RA/4 is contraindicated in patients whom the physician determines are not candidates for surgical procedures due to physical or mental conditions. ▪ Exclusion criteria: This device should be kept out of the reach of children. Safety has not been established for pregnancy, patients under the age of 18, patients with suspected or real heart problems, or patients who have implanted electrical devices or epilepsy. The long-term effects of electrical stimulation of the diaphragm are unknown. This device is electrically powered and may produce tissue damage or electrical hazard if improperly used. The system may be affected by excessive moisture, severe mechanical shock, diathermy, electro cauterization, and radiation therapy. Implanted patients should not be connected to high-frequency surgical equipment or subjected to magnetic resonance imaging (MRI). Care should be taken to avoid operation of this device in close proximity to shortwave or microwave therapy equipment. Discontinue use of this device if skin in the implant area becomes swollen, infected, or inflamed or if there are skin eruptions such as phlebitis, thrombophlebitis, or varicose veins. Adverse events related to the system include capnothorax, equipment failure leading to loss of breathing, infection, airway compromise, spasms, pain or discomfort with stimulation, and difficulty eating. Patients must have a mechanical ventilator available at all times
2024-01-16 | HTA23071 | Blood Biomarker for mild Traumatic Brain (I-Stat-TBI) | Abbott Laboratories GMBH | Abbott Point of Care | USA | Diagnostic Test | Diagnostic | 1. The technology should be performed by authorized healthcare laboratory. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to DOH upon request. Population, setting and intended user for Technology “Blood Biomarker for mild Traumatic Brain Injury” ▪ Population/ Intended User; – patient suspected mild traumatic brain injury (Glasgow Coma Scale 13-15) within 12 hours of injury. ▪ To be performed by: – Medical Lab Technology ▪ Clinical Setting: – Hospitals, Emergency Department ▪ Condition of use: – 18 years of age or older – Glasgow coma scale 13-15 as the Negative Predictive Value (NPV) is 99.3% .
2024-01-09 | HTA23043 | Cochlear Implant System | AbuDhabi University | NEUBIO AG | India | Device | Therapeutic | 1. Cochlear Implant System/ BOLD I should be performed by (ENT) specializes. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to DOH upon request Population, setting and intended user for Technology “” ▪ Population/ Intended User; – Bilateral profound sensorineural hearing loss – Limited benefit from amplification/hearing aid. ▪ To be performed by: – By Ear, Nose, and Throat doctor (ENT) specializes. ▪ Clinical Setting: – Hospitals, special ENT surgery centers ▪ Condition of use: – The BOLD Cochlear Implant System aims to restore a sensation of hearing in individuals aged 12 months and older. ▪ Exclusion criteria: – The BOLD CI system is contraindicated for individuals with conditions such as deafness. – middle ear infections. – cochlear ossification preventing electrode insertion. – absence of cochlea
2024-01-02 | HTA23064 | My Care Oncology Busulfan Assay Kit | Al Nawras Medi-Lab Supplies | SALADAX | USA | Diagnostic Test | Diagnostic | Population, setting and intended user for Technology “My Care Oncology Busulfan.” ▪ Population/ Intended User; – A patient with multiple myeloma who is receiving busulfan conditioning prior to HSCT (haematopoietic stem cell transplantation). – A patient with sickle cell disease who is receiving busulfan conditioning prior to HSCT. – A patient with aplastic anemia who is receiving busulfan conditioning prior to HSCT. – A patient with thalassemia who is receiving busulfan conditioning prior. ▪ To be performed by: – Medical Lab Technology, ▪ Clinical Setting: – Hospitals, Medical laboratories. ▪ Condition of use: – NA ▪ Exclusion criteria: – pregnant women. – Terminal illness
2023-12-20 | HTA23066 | Quidel Triage PlGF Test | Gulf Med Medicines LLC | QuidelOrtho Corporation | USA | Diagnostic Test | Diagnostic | 1. Approval of Technology Quidel Triage PlGF Test (98800EU) and Quidel Triage PlGF Control 1 (98813EU) and Control 2 (98814EU) as an aid test alongside with other clinical information to help diagnose preterm pre-eclampsia in women who are between 20 weeks and 35 weeks pregnant with signs and symptoms of pre-eclampsia. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Provision of regular updates and reports about the product to DOH upon request. 4. Communicate the need for screening strategy for pre-eclampsia to ADPHC utilizing the new technologies such as this test “Quidel Triage PlGF Test”. Population, setting and intended user for Technology “Quidel Triage PlGF Test” ▪ Population/ Intended User; – The test is used with other clinical information to help diagnose preterm pre-eclampsia in women who are between 20 weeks and 35 weeks pregnant with signs and symptoms of preeclampsia. ▪ To be ordered by: – By Obstetrics and gynecology ▪ Clinical Setting: – Healthcare Facilities ▪ Condition of use: – As per the manufacturer instructions ▪ Exclusion criteria: – Other conflicting medical issues
2023-12-06 | HTA23061 | Thermogenesis AXP II system | Al Nawras Medi-Lab Supplies | ThermoGenesis | USA | Device | Therapeutic | Population, setting and intended user for Technology “Thermogenesis AXP II system” ▪ Population/ Intended User; – Patient elegable for stem cell therapy – Parents opting for vord blood banking ▪ To be performed by: – Trained Healthcare professionals ▪ Clinical Setting: – Hospitals – Clean environment ▪ Condition of use: – Operation should be conducted by individuals who have received training and certification specific to the AXP® II System – The system should be used in a controlled, clean, and sterile laboratory environment to maintain the integrity of the cord blood and prevent contamination. – Regular maintenance and calibration of the system are necessary to ensure its optimal performance and accuracy. – Users must adhere to the regulatory standards and guidelines set by health authorities. – The system should be used for its intended purpose of processing cord blood for stem cell concentration, and not for other unapproved applications. ▪ Exclusion criteria: – Cord blood that is visibly contaminated or of poor quality, which might not yield viable stem cells upon processing. – The system may require a minimum volume of cord blood to operate effectively. Samples below this threshold might be excluded. – If the blood donor (typically the newborn or the mother) has certain infectious diseases, this might preclude the use of the cord blood due to safety concerns. – Cord blood units that have exceeded their viable storage period might not be suitable for processing
2023-11-27 | HTA23056 | EO Patch Wearable Insulin Pump | Gulf Drug LLC | EO FLOW | South Korea | Device | Therapeutic | Population, setting and intended user for Technology “EO Patch Wearable Insulin Pump” ▪ Population/ Intended User; – Type 1 Diabetes Mellitus ▪ To be performed by: – Healthcare professionals – Trained individual with diabetes – Trained parents whose children have diabetes ▪ Clinical Setting: – Hospitals – Clean environment ▪ Condition of use: – The system is intended to be used by insulin-dependent persons with diabetes mellitus, aged ≥ 6 years old ▪ Exclusion criteria: – People who regularly require less than 0.1 U/h of basal insulin – People who are not able to be in regular contact with their healthcare professional. – People who do not understand what is required for insulin pump therapy or who are not able to follow the instructions for use of the system. – People with skin that does not tolerate adhesive pads. – People who are not able to notice alarms because of physical limitations.
2023-11-24 | HTA23039 | High Intensity Focused Ultrasound (HIFU) | Cleveland Clinic Abu Dhabi | Cleveland Clinic Abu Dhabi | france | Device | Diagnostic | Population, setting and intended user for Technology “HIFU” ▪ Population/ Intended User; – Patients who have not yet received any treatment for prostate cancer. – HIFU is particularly recommended for radical treatment with patients who have: • A localized stage t1 or t2 cancer. • A Gleason score of 7 or less. • A desire to maintain maximum quality of life after treatment. ▪ To be performed by: – By urologist. ▪ Clinical Setting: – Hospitals, special urology centers ▪ Condition of use: – prostate cancer. ▪ Exclusion criteria: – Anal or rectal surgery that prevents insertion of the probe. – Artificial sphincter, penile implants and intra-prostatic implants. – Latex allergy (due to the composition of the balloon surrounding the treatment probe). – Inflammatory Bowel Disease (IBD), previous significant rectal surgeries, urethral strictures, active infections, and anatomical inability to insert the transrectal probe
2023-11-03 | HTA22062 | WATCH PAT ONE, DEVICE | Medworks Medical Equipment LLC | ITAMAR MEDICAL LTD | ISRAEL | Device | Therapeutic | ▪ Population/ Intended User; – Patient with Sleep Apnea/ patients suspected to have sleep related breathing disorders. ▪ To be prescribed by: – By Healthcare Professionals. ▪ Clinical Setting: – NA ▪ Condition of use: – As per the device manual ▪ Exclusion criteria: – WATCH PAT ONE can be used only for the screening/diagnosis of moderate to severe OSA without other comorbidities like cardiovascular disease, stroke, or other conflicting medical issues
2023-10-16 | HTA23050 | SleepUp | Propharma Medical Supplies | SleepUp Tecnologia em Saúde Ltda | Brazil | Device | Therapeutic | 1. The approval is limited to one year only, the applicant shall return with further evidence on the platform efficacy for reevaluation. 2. Recommended a medical assessment should be done before referral to SleepUp during pregnancy and in people with comorbidities. 3. SleepUp is not intended to provide diagnostics and does not replace a doctor's consultation nor prescribe medication. 4. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 5. Provision of regular updates and reports about the product to DOH upon request. Population/ Intended User; – The treatment of insomnia and subjective sleep complaints in adults. ▪ To be prescribed by: – NA ▪ Clinical Setting: – NA ▪ Condition of use: – As per Healthcare professional guidance ▪ Exclusion criteria: – Other conflicting medical issues
2023-10-02 | HTA23044 | DropleX KRAS Mutation Test v2 | Al Nawras Medi-Lab Supplies | Gencurix Inc. | Republic of Korea | Product/Kit | Diagnostic | Population/ Intended User; – patients currently getting treatment. – patient with new cancer incidents, who need to be screened prior to prescribed for treatment. To be performed by: – Medical lab technologist. Clinical Setting: – Hospitals, special laboratory. Condition of use: – The test helps identify CRC and NSCLC patients who are eligible for targeted treatment. – Monitoring targeted therapy for cancer mutation development. Exclusion criteria: – RAS mutation; MEK inhibitors; Pancreatic ductal adenocarcinoma; Prognosis; Quantitative PCR; Solid pancreatic masses.
2023-09-28 | HTA23045 | Endosafe®-PTS™ system & Endosafe®-PTS™ cartridges | Al Nawras Medi-Lab Supplies | Charles River | USA | Device | Diagnostic | Population/ Intended User; – NA ▪ To be performed by: – By medical lab technology and clinical pharmacist ▪ Clinical Setting: – Hospitals, enviromenal lab , pharmacy, in-centre and community dialysis units (CDUs) thatprovide hemodialysis (HD) and/or hemodialfiltration(HDF). ▪ Condition of use: – All injectable pharmaceuticals. – Implantable medical devices. – Water dialysis unit – Surgical Instruments, Cleanroom and Gloves. – Cell therapy. ▪ Exclusion criteria: – This test is not meant to identify endotoxins for patient liquid samples or to make a medical diagnosis.
2023-09-21 | HTA23040 | TRANSCUTANEOUS BILIRUBIN METER (BILICARE) | Hessa Medical Equipments Co. LLC | Gerium Medical | Israel | Diagnostic Test | Diagnostic | Population/ Intended User; – For screening infants for hyperbilirubinemia. ▪ To be performed by: – By healthcare professionals. ▪ Clinical Setting: – Healthcare facilities. ▪ Condition of use: – As per Doctor recommendation. ▪ Exclusion criteria: – Other conflicting medical issues.
2023-09-21 | HTA23022 | Second Generation Fecal Immunochemical Test (FIT) | Pharma Plus Drug Store LLC | BioLabs | USA | Product/Kit | Diagnostic | ▪ Population/ Intended User. – aged between 40-65 years. – Patient with hemoglobin in feces. ▪ To be performed by: – By medical professional . ▪ Clinical Setting: – Hospitals, special Laboratory centers ▪ Condition of use: – A personal history of colorectal cancer or certain types of polyps – A family history of colorectal cancer – A personal history of inflammatory bowel disease (ulcerative colitis or Crohn’s disease) – A confirmed or suspected hereditary colorectal cancer syndrome, such as familial adenomatous polyposis (FAP) or Lynch. – syndrome (hereditary non-polyposis colon cancer or HNPCC). – A personal history of getting radiation to the abdomen (belly) or pelvic area to treat a prior cancer. ▪ Exclusion criteria: – Menstrual bleeding. – Bleeding hemorrhoids. – Constipation bleeding. – Urinary bleeding. Alcohol and certain medications such as aspirin, indomethacin, reserpine, phenylbutazone, corticosteroid and nonsteroidal anti-inflammatory drugs may cause gastrointestinal irritation and subsequent bleeding in some patients
2023-09-18 | HTA23032 | Contrast Enhanced Mammogram (CEM) Service | AHS- SEHA | General Electric Healthcare | Cleveland Clinic Abu Dhabi | Technique/Procedure | Diagnostic | Population/ Intended User. – patients with dense breasts. – used to evaluate breast symptoms or abnormalities seen on mammogram. – to assess the extent of cancer in women with newly diagnosed breast cancer and to monitor response to neoadjuvant chemotherapy prior to surgery. ▪ To be performed by: – Specialist mammographer/Radiographer. ▪ Clinical Setting: – Hospitals and special mammogram centers ▪ Condition of use: – who has received an inconclusive result after standard mammogram being referred to MRI – The Quality of Tumor Size Assessment ▪ Exclusion criteria: – Do not perform CEM exam for patients with implants
2023-08-11 | HTA23018 | Elecsys PlGF & Elecsys sFlt-1 | Roche Diagnostics | Roche Diagnostics | Germany | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – Obstetrics and gynaecology ▪ To be performed by: – Trained Laboratory staff ▪ Clinical Setting: – Hospitals – Laboratory setting ▪ Condition of use: – As per manufacturer guidelines ▪ Exclusion criteria: – Exclusion criteria may include specific medical conditions or comorbidities that could interfere with the test results or pose a risk to the individual. For example, if the test involves drawing blood, individuals with bleeding disorders or those on anticoagulant medications may be excluded. – Individuals who are taking certain medications or have undergone specific interventions that could affect the test results or its interpretation
2023-08-11 | HTA23033 | THYROID ABLATION By laser | Mediclinic Hospital(Al Noor) | El.En Electronic Engineering Spa | Italy | Device | Therapeutic | Population/ Intended User; ➢ Thyroid Laser ablation can be used in the following patients with proven benign thyroid nodules [usually will need 2 benign biopsies]: – Neck compressive symptoms such as dysphagia, dyspnea or neck discomfort due to large thyroid nodule(s) – Cosmetic concerns about the size and/or appearance of the goiter – Autonomous thyroid nodules causing subclinical or overt hyperthyroidism ▪ To be performed by: – The laser ablation therapy performed by interventional radiologists. ▪ Clinical Setting: – Outpatient setting; Hospitals, special surgery centers. ▪ Condition of use: ➢ LA is not first-line treatment for malignant or suspicious thyroid lesions, but has been used in the following conditions: – Papillary thyroid microcarcinoma where patient is unwilling or unfit for surgery. – Surgically Unresectable thyroid cancer. – Cervical lymph-node recurrence of differentiated thyroid cancer. ➢ it is usually requiring two benign FNA (fine needle aspiration) before implementing the procedure). ▪ Exclusion criteria: ➢ LA is not suitable for patients with: – Untreated clotting disorders – Thyroid nodules with high suspicion features for malignancy on ultrasound – Abnormal vocal cord function
2023-07-28 | HTA23027 | Z50 Hematology Analyzer | Al Bahar Al Azraq Drug Store | Zybio INC | China | Device | Diagnostic | Population/ Intended User; – To analyses blood cells. ▪ To be performed by: – By medical lab technician/technologist. ▪ Clinical Setting: – Laboratories. ▪ Condition of use: – As per manufacturer instruction. ▪ Exclusion criteria: – Other conflicting medical issues
2023-07-28 | HTA23028 | Chemistry Analyzer EXC200 | Al Bahar Al Azraq Drug Store | Zybio INC | China | Device | Diagnostic | Population/ Intended User; – to analyze biological samples such as blood, serum or plasma. ▪ To be performed by: – By medical lab technician/technologist. ▪ Clinical Setting: – Hospitals, clinics, laboratories. ▪ Condition of use: – As per manufacture instruction. ▪ Exclusion criteria: – Other conflicting medical issues
2023-07-28 | HTA23030 | EarWell® (nonsurgical infant ear deformity correction device) | MEDICAL IDEA MEDICAL SURGICAL ARTICLES & REQUISITES TRADING L.L.C | Becon Medical | USA | Device | Therapeutic | Population/ Intended User; – Infant with ear deformities at early age (up to three weeks) ▪ To be performed by: – By Healthcare Professional ▪ Clinical Setting: – At doctors’ clinic. ▪ Condition of use: – As per the manufacturer instruction and doctor prescription. ▪ Exclusion criteria: – Other conflicting medical issue
2023-07-25 | HTA23015 | Bio-Rad QXDx AutoDG ddPCR System | Al Nawras Medi-Lab Supplies | BIORAD | USA | Device | Diagnostic | ▪ Population/ Intended User; – Oncology, infectious disease, genetic disorders, personalized medicine and transplantation patients. ▪ To be performed by: – be used by healthcare professionals. ▪ By Clinical Setting: – Hospitals and specialized clinics. ▪ Condition of use: – Patients who are suspected or diagnosed with a specific disease or condition that can be detected or monitored using the targeted assays on the ddPCR systems. ▪ Inclusion criteria: – Patients with specific genetic mutations, biomarkers, or genetic variations of interest that are known to be associated with the disease or condition being diagnosed or monitored. These targets should have clinical relevance and established evidence supporting their use for diagnostic or prognostic purposes. ▪ Exclusion criteria: – Patients for whom there is insufficient sample material available for testing may be excluded. – Patients for whom the diagnostic context or the intended use of ddPCR is not clinically relevant or appropriate. – Patients who do not have the specific genetic mutations, biomarkers, or genetic variations targeted by the ddPCR assays. – Patients for whom the use of ddPCR does not offer clinically significant benefits or contribute to decision-making in their specific diagnostic or treatment context.
2023-06-26 | HTA23038 | Omnipod DASH® Insulin Management System | Insulet Middle East & Developing Markets as a branch office of Insulet Netherlands B.V. | Insulet Corporation | USA | System/Software/ AI/App | Therapeutic | ▪ Population/ Intended User; – Diabetes Mellitus ▪ To be performed by: – Healthcare professionals – Trained individual with diabetes – Trained parents whose children have diabetes ▪ Clinical Setting: – Hospitals – Clean environment ▪ Condition of use: – The system is intended to be used by insulin-dependent persons with diabetes mellitus, aged ≥ 6 years old ▪ Exclusion criteria: – The micropump system should not be used by children under 6 years of age – People who regularly require less than 0.1 U/h of basal insulin – People who are not able to be in regular contact with their healthcare professional. – People who do not understand what is required for insulin pump therapy or who are not able to follow the instructions for use of the system. – People with skin that does not tolerate adhesive pads. – People who are not able to notice alarms because of physical limitations.
2023-06-13 | HTA23026 | DropleX EGFR Mutation Test v2 | Al Nawras Medi-Lab Supplies | Gencurix Inc. | Republic of Korea | Diagnostic Test | Diagnostic | Population/ Intended User; – Prevalent major cancer patients currently getting treatment, as well as patient with new cancer incidents, who need to be screened prior to prescribed for treatment. ▪ To be performed by: – be used by healthcare professionals. ▪ By Clinical Setting: – Hospitals and specialized clinics (oncology) ▪ Condition of use: – EGFR mutation testing is ordered after the patient has been diagnosed with non-small cell lung cancer, especially adenocarcinoma. ▪ Exclusion criteria: – These tests are not helpful for identifying patients with lung cancer who may benefit from EGFRtargeted tyrosine kinase inhibitor therapy
2023-06-09 | HTA23031 | Transcatheter Bicaval Valve System (TricValv) | Cleveland Clinic Abu Dhabi | P+F Products and Features GmbH | Germany | Technique/Procedure | Therapeutic | ▪ Population/ Intended User; – patients with symptomatic severe tricuspid regurgitation. – The patient must be 18 years of age or older. ▪ To be performed by: – Interventional cardiologist, and a cardiologist. ▪ Clinical Setting: – Hospitals, Special cardiology centers. ▪ Condition of use: – The Medical Device is indicated for treatment of hemodynamically significant tricuspid regurgitation (TR) and caval reflux in patients who have been determined to be at extreme risk for tricuspid valve surgery as determined by multidisciplinary heart team experienced in the evaluation and treatment of heart failure and tricuspid valve disease. ▪ Exclusion criteria: The TricValve Transcatheter Bicaval Valves System is contraindicated in patients who have any of the following conditions: – Evolutionary or recent stroke. – Cerebrovascular accident (CVA) evolutionary or recent. – Recent myocardial infarction (<30 days). – Known hypersensitivity, allergy or contraindication to device’s components, e.g. nitinol, bovine pericardium etc. – Known hypersensitivity to vitamin K antagonists, heparin and other oral anticoagulants, or sensitivity to contrast medium that cannot be adequately premedicated. – Sepsis, including active endocarditis. – Thrombosis of the lower venous system or vena cava filter. – Contraindication against a transesophageal echo (TEE) during the procedure. – Creatinine clearance <20 ml / min. – Vascular conditions (e.g. stenosis, tortuosity) that make insertion and endovascular access impossible to the upper and lower vena cava. – Bleeding diathesis or coagulopathy or patient refusing blood transfusion. – Active gastritis or peptic ulcer. – Pregnancy. – inhibitor therapy
2023-06-06 | HTA23012 | Alio 1.2 | Alio Inc. | Alio, Inc. | USA | System/Software/ AI/App | Diagnostic | ▪ Population/ Intended User; – Patients with Chronic Kidney Disease (CKD) ▪ To be performed by: – By Healthcare Professionals ▪ Clinical Setting: – Clinical setup process ▪ Condition of use: – As per the manufacturer instruction ▪ Inclusion criteria: – This product applies patients who are 18 years of age or older with End Stage Kidney Disease (ESKD) who are undergoing dialysis via an arm-based fistula or graft ▪ Exclusion criteria: – Skin surface issues located at the Alio placement site(s) (AVG/ AVF) that have the potential to cause subject discomfort or interfere with data collection
2023-05-26 | HTA22030 | Guardant360® | Guardant Health | Guardant Health Inc | USA | Diagnostic Test | Diagnostic | Guardant360 is not indicated for the diagnosis or minimal residual disease detection in cancer patients. ▪ Population/ Intended User; - Cancer patients. Guardant360 is indicated for use only in patients with an established diagnosis of an advanced solid tumor and test requests without such confirmation from the physician are canceled by the clinical laboratory. - Guardant360® can be applied for the monitoring and aftercare of a cancer therapy as well - Indicated for: ▪ Advanced solid tumors ▪ Before first-line therapy or at progression ▪ To be ordered by: – Oncologist. ▪ Clinical Setting: – Authorized Laboratories (The test analyzed abroad in US in the laboratory of Guardant Health) ▪ Condition of use: – The specimen collection & shipping shall follow the international standards & manufacturer guidelines. ▪ Exclusion criteria: – Not indicated for: ▪ Hematologic malignancies ▪ Early stage cancers ▪ When disease is stable or responding to therapy
2023-05-03 | HTA23002 | Q-Suite (Pantheon and AI Modules) | Qritive Pte Ltd. | Qritive Pte Ltd | Singapore | System/Software/ AI/App | Diagnostic | conditional approval on use the Q-Suite (Pantheon and AI Modules) in a clinical trial setting to validate the model's outcome ▪ Population/ Intended User; – For the detection of cancer cases (Prostate, Colon and Lymph Node AI & IHC) ▪ To be performed by: – In Pantheon users are restricted to four user groups based on different user requirements. These are configurable using role-based access implemented in Pantheon. These user groups are: • Pathologists • Pathology technicians • Pathology lab administrators • Other clinicians ▪ Clinical Setting: – Healthcare facilities and laboratories, telepathology. ▪ Condition of use: – As per manufacturer instructions. ▪ Exclusion criteria: – NA.
2023-04-05 | HTA23005 | Elecsys Troponin T hs, Elecsys Troponin T hs STAT (0/1 hour algorithm) | Roche Diagnostics | Roche Diagnostics | Germany | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; ▪ To be performed by: – By Healthcare Professionals ▪ Clinical Setting: – Laboratories ▪ Condition of use: – Patients presenting to the ER with chest pain ▪ Exclusion criteria: – Other conflicting medical issues
2023-03-31 | HTA22064 | Stereo-electroencephalography (SEEG) | Cleveland Clinic Abu Dhabi | DIXI medical | France | Device | Therapeutic | ▪ Population/ Intended User; – Patients with epilepsy whose seizures do not successfully respond to antiseizure medication therapy is considered to have drug-resistant epilepsy (DRE). – Have focal epilepsy and seizures (complex partial seizures) that do not respond to two anticonvulsant medications or medical treatment. – Are a potential candidate for epilepsy surgery. – If the origin sites of the patient seizures can’t be found with other tests. ▪ To be performed by: – By neurosurgeons ▪ Clinical Setting: – Hospitals, special surgery centers, operating room/theatre ▪ Condition of use: – As per the followed standards for epilepsy patients. ▪ Exclusion criteria: – Patients who are not eligible for brain surgery. – Other conflicting health issue that prevent from doing the surgery.
2023-03-27 | HTA22052 | NT-proBNP | Roche Diagnostic Middle East | Roche Diagnostics | Germany | Diagnostic Test | Diagnostic | NT-proBNP is a standardized diagnostic tool for suspected heart failure cases. Using NT-proBNP as an aid diagnostic test for diabetic patients type 2 for the diagnosis of heart failure in conjunction with the clinical information, examination and other data available to the physician. ▪ Population/ Intended User; – NT-proBNP is indicated as an aid in the diagnosis of individuals suspected of having congestive heart failure and detection of mild forms of cardiac dysfunction. ▪ To be performed by: – To be ordered by healthcare professional. ▪ Clinical Setting: – Authorized laboratories ▪ Condition of use: – The test ordered for DM2 patients should be considered in conjunction with the clinical information and other data available to the physician for high risk population. ▪ Exclusion criteria: – Other conflicting medical issues
2023-03-10 | HTA22053 | DxPG80.Lab | Balsam Health Services | Progastrin Manufacturing | France | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; *hPG80 detects more than 16 types of cancer - Brain tumors - All Cancers of Head and Neck - Skin cancer - Lung cancer - Liver cancer - Pancreatic cancer - Stomach cancer - Colon cancer - Rectal cancer - Prostate cancer - Uterine cancer - Cervical cancer - Neuroendocrine tumors - Ovarian cancer - Kidney cancer - Breast cancer ▪ To be order by: – By ONCOLOGIST ▪ Clinical Setting: – Hospitals, Medical centers. ▪ Condition of use: – As per Manufacturer instruction. ▪ Exclusion criteria: - Patient: Not for children (under 18 years old) or for pregnant women. - Professional users of the test: Label warning: Caution, consult instructions for use, contains human blood or plasma derivatives, contains biological material of human origin, keep in dry conditions, do not use if package is damaged, Warning: Irritant.
2023-03-06 | HTA23004 | Abbott ID NOW™ | Abbott | Abbott Diagnostics Scarborough, Inc. | USA | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – Using IDNOW influenza A and B / Strep A /RSV for symptomatic patients. ▪ To be performed by: – By Healthcare Professionals ▪ Clinical Setting: – Point of care ▪ Condition of use: – As per the manufacturer instruction ▪ Exclusion criteria: – Other conflicting medical issues
2023-03-03 | HTA22058 | CFX Opus Dx Real-Time PCR Detection Systems for In Vitro Diagnostics (IVD) | Al Nawras Medi-Lab Supplies | BIORAD | USA | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – For In vitro diagnostic test ▪ To be performed by: – By Lab technicians ▪ Clinical Setting: – clinical laboratory. ▪ Condition of use: – As per the manufacturer guidance manual ▪ Exclusion criteria: – As per the manufacturer guidance manual
2023-02-28 | HTA22059 | Tablo Hemodialysis System | Global Medical Supply Chain | Outset Medical, Inc. | USA | Device | Therapeutic | ▪ Population/ Intended User; – Patients with acute and/or chronic renal failure. ▪ To be performed by: – By trained healthcare professionals. ▪ Clinical Setting: – Acute or chronic care facility. – The Tablo Hemodialysis System is also indicated for use in the home. ▪ Condition of use: – Treatments must be administered under physician’s prescription and observed by a trained individual who is considered competent in the use of the device. ▪ Exclusion criteria: – If the patient has condition that prevent him/her from conducting dialysis.
2023-02-07 | HTA22044 | Dry Eye treatment-Rexoneye | Mediclinic Al Noor | Resono | Italy | Device | Therapeutic | conditional approval of using this technology and The applicant shall refer back to DOH after 6 Months from date of approval with more evidence on the clinical and cost effectiveness of the treatment for re-evaluation. ▪ Population/ Intended User; – Patients who suffer from dry eye syndromes. ▪ To be performed by: – By Ophthalmologists. ▪ Clinical Setting: – Hospitals, ophthalmology clinic. ▪ Condition of use: – Clinical setting. ▪ Exclusion criteria: Rexon Device could not be used on the following – Pregnant women – Patients carrying active implantable devices – Neonatal or pediatric Patients – Patient who undergo ocular surgeries past month
2023-02-03 | HTA22040 | Holter Connect Platform based on LifeSignals ECG Remote Monitoring Patch Platform. | les source medical supplies LLC | LifeSignal, Inc | USA | Device | Diagnostic | ▪ Population/ Intended User; – Non-critical, adult population, who are 18 years of age or older ▪ To be performed by: – By healthcare professionals ▪ Clinical Setting: – Home and Clinical Facilities ▪ Condition of use: – As per the manufacturer instructions ▪ Exclusion criteria: – The Biosensor is not intended for use on critical care patients. – The Biosensor is not intended for use on patients with any active implantable devices, such as defibrillators or pacemakers – In case the patient has skin condition or allergic to adhesives or electrode hydrogels.
2022-12-29 | HTA22043 | Exsurgo – Axon | Richard Little | Exsurgo Ltd | New Zealand | Device | Therapeutic | Approve with limitation Population/ Intended User; – Management of Chronic pain from any neurological or musculoskeletal cause, of which examples could be, post herpetic (or post-shingles) neuralgia, complex regional pain syndrome (CRPS), lower back pain, neck pain, major joint pain, spinal cord injury, phantom limb pain, brachial plexus injury related pain, traumatic peripheral nerve injury pain and post-cancer treatment pain, it may have potential on Mental health as well such as anxiety. ▪ To be performed by: – By certified medical professional. ▪ Clinical Setting: – It is a prescription device that could be used in both clinic and home. ▪ Condition of use: – As per the manufacturer instructions for use. ▪ Exclusion criteria: – Other conflicting medical issues
2022-12-09 | HTA22045 | Tata MD Check RT-PCR Fast 3Gene | Tata Medical And Diagnostics Ltd | Tata Medical & Diagnostics Ltd., | India | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – Using Tata MD Check RT-PCR Fast 3Gene for covid 19 testing. ▪ To be performed by: – By Healthcare Professionals ▪ Clinical Setting: – Point of care ▪ Condition of use: – As per the manufacturer instruction ▪ Exclusion criteria: – Other conflicting medical issues
2022-11-22 | HTA22055 | Xpert Xpress SARS-CoV-2/Flu/RSV | AL Zahrawi medical group | Cepheid | USA | Diagnostic Test | Diagnostic | Population/ Intended User; – Using Xpert Xpress SARS-CoV-2/Flu/RSV for symptomatic patients. ▪ To be performed by: – By Healthcare Professionals ▪ Clinical Setting: – Point of care ▪ Condition of use: – As per the manufacturer instruction ▪ Exclusion criteria: – Other conflicting medical issues
2022-11-14 | HTA22033 | COVID-19 PCR KIT-ASTRAGENE | Dr. Palat Krishna Menon | ASTRAGENE FZ LLC | DUBAI, United Arab Emirates | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – Using COVID-19 PCR KIT-ASTRAGENE for symptomatic patients. ▪ To be performed by: – By Healthcare Professionals ▪ Clinical Setting: – Point of care ▪ Condition of use: – As per the manufacturer instruction ▪ Exclusion criteria: – Other conflicting medical issues
2022-10-31 | HTA22047 | Y-90 RADIOEMBOLISATION | Cleveland Clinic Abu Dhabi | Sirtex Medical Pty. Ltd | Australia | Technique/Procedure | Therapeutic | Population/ Intended User; Radioembolization is a therapy used to treat advanced-stage liver cancer that has either started in the liver or spread to the liver from another part of the body (metastatic). Treatment of primary liver tumors (Intrahepatic Cholangiocarcinoma, as well as other liver metastatic disease (NETs,Uveal Melanoma, Breast Cancer, …) is off label. To be performed by: Radioembolization is an outpatient treatment performed by an interventional radiologist with training in interventional oncology. The treatment is commonly performed with sedation and local anesthesia. Clinical Setting: Hospitals Condition of use: Y-90 to treat tumors that were initially formed in the liver or have spread (or metastasized) to the liver from another part of the body. It is a palliative treatment. Radioembolization procedures are typically performed using x-ray guidance. Exclusion criteria: Radioembolization is not recommended in cases of severe liver or kidney dysfunction, abnormal blood clotting or a blockage of the bile ducts.
2022-10-25 | HTA22037 | COVID-19 (SARS-CoV-2) Nucleic Acid Test Kit (Freeze-dried) | EKRAMY MOHAMED MOHAMED | Wuhan EasyDiagnosis Biomedicine Co., Ltd. | China | Diagnostic Test | Diagnostic | •Population/ Intended User; Using COVID-19 (SARS-CoV-2) Nucleic Acid Test Kit (Freeze-dried) for covid 19 testing for symptomatic patients.( suspected cases) •To be performed by: By Healthcare Professionals •Clinical Setting: Point of care •Condition of use: As per the manufacturer instruction. •Exclusion criteria: Other conflicting medical issues
2022-10-25 | HTA22048 | SARS-CoV-2 Nucleic Acid Rapid Detection Kit- Fluorescence RT-PCR system | Hangzhou Bioer Technology Co., Ltd. | Hangzhou Bioer Technology Co., Ltd. | China | Diagnostic Test | Diagnostic | •Population/ Intended User; Using SARS CoV 2 Nucleic Acid Rapid Detection Kit (Fluorescence RT PCR) for covid 19 testing for symptomatic patients.( suspected cases) •To be performed by: By Healthcare Professionals •Clinical Setting: Point of care •Condition of use: As per the manufacturer instruction •Exclusion criteria: Other conflicting medical issues
2022-09-30 | HTA22031 | Allplex SARS COV2/FLU A/FLU B/ RSV Assay | Lubna Bdour | Seegene | SOUTH KOREA | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – For SARS-COV-2, Flu A and Flu B, and RSV testing in Symptomatic cases. ▪ To be performed by: – By healthcare professional. ▪ Clinical Setting: – Point of care. ▪ Condition of use: – For COVID 19 suspected cases, Flu A and Flu B, and RSV testing ▪ Exclusion criteria: – Other conflicting medical issues
2022-09-20 | HTA22034 | Cryoflex CRYO-S Painless | Reda Tolba, MD | Metrum | Poland | Device | Therapeutic | ▪ Population/ Intended User; – Patients with chronic pain – Best patients for cryoanalgesia are those with chronic pain at the level of over 5 in visual alogue scale (VAS) – Clinical application: – Knee pain prior or secondary to infrapatellar branch of the saphenous nerve, anterior femoral cutaneous nerve, genicular nerve after knee arthroplasty or total knee replacement – Various types of peripheral neuropathy – Diabetic neuropathy – Low back pain and lower extremity pain secondary to lumbar facet joint pathology, pseudosciatica – Phantom pain – Morton’s neuroma, heel pain – Sacroiliac joint pain, cluneal neuralgia, obturator neuritis – Chest wall pain with multiple conditions including post-thoracotomy neuromas, persistent pain after rib fractures and post herpetic neuralgia in thoracic distribution – Abdominal and pelvic pain secondary to ilioinguinal, iliohypogastric, genitofemoral, subgastric neuralgia, pudendal neuralgia – Craniofacial pain secondary to trigeminal neuralgia, posterior auricular neuralgia, and glossopharyngeal neuralgia – Upper extremity pain secondary to suprascapular neuritis and other conditions of peripheral neuritis ▪ To be performed by: – Recommended for Pain Management Specialists: – Anesthesiologists – Neurosurgeons – Orthopedists – Spine Surgeons – Interventional Radiologist ▪ Clinical Setting: – Hospitals and specialized clinics. ▪ Condition of use: – Procedure recommended especially when traditional methods prove to be unsatisfactory. ▪ Exclusion criteria: – Other conflicting medical issues
2022-09-12 | HTA22029 | Emsella (HIFEM) Pelvic Floor Strengthening Device | Dr.Tracey Sims | BTL | Czech Republic | Device | Therapeutic | Population/ Intended User; ▪ Electromagnetic stimulation of pelvic floor musculature for the purpose of rehabilitation of weak pelvic muscles ▪ To be performed by: – Healthcare professionals ▪ Clinical Setting: – Hospitals – Clean environment ▪ Condition of use: – The system is intended to be used for restoration of neuromuscular control for the treatment of male and female urinary incontinence – Stress urinary incontinence, urge incontinence, or mixed urinary incontinence as determined by Healthcare professionals – Age greater than 18 – Body mass index (BMI) < 37 kg/m2 4 ▪ Exclusion criteria: – Patient with any significant pelvic organ prolapse; stage III or greater – Patient with untreated malignancy – Patient who is pregnant, planning to get pregnant or within 3 months postpartum – Patient with implant or IUD containing metal – Patient has a pacemaker
2022-08-12 | HTA22028 | seca mBCA 554 | Hasan Abdallah | Seca gmbh & Co. Kg | Germany | Device | Diagnostic | Population/ Intended User; – Any patients of any age groups except babies ▪ To be performed by: – By Healthcare professional ▪ Clinical Setting: – Clinical settings ▪ Condition of use: – As per Manufacturer instructions ▪ Exclusion criteria: – Bioimpedance Analysis may not be performed on persons with: – Clinical use of BIA: should not be used for the below mentioned subjects as there no existing accurate, validated and proven BIA algorithms/equations for interpretation/estimation of the Body masses. o Extremes of BMI ranges – (<16 or >34 kg/m2) o In severely malnourished and anorexia nervosa patients (BMI <16 kg/m2), results of BIA can dangerously impact the management of Refeeding syndrome of such subjects. o Medical conditions which change serum electrolytes, hematocrit and blood flow, o All Subjects with abnormal hydration and fluid statuses. o Subjects who are undergoing Chemotherapy & Radiation therapy. o Neuromuscular diseases (e.g. Duchenne muscular dystrophy, post-traumatic paraplegia/hemiplegia) o Patients having HIV-infection, and liver cirrhosis o Patients on hemodialysis, peritoneal dialysis, and chronic renal failure. o Patients with acromegaly. o Pregnant women. o Elderly o Patients having Implanted devices and metallic devices
2022-08-12 | HTA21013 | Dexcom G6 System | Yaser Hosny | DEXCOM, INC. | USA | Device | Diagnostic | ▪ Population/ Intended User; – Type 1 and Type 2 DM (uncontrolled) – Gestational Diabetes (pregnancy) – Diabetic with High A1C, on multiple daily doses of Insulin ▪ To be performed by: – Self-Used ▪ Clinical Setting: – NA ▪ Condition of use: – As per the manufacturer instructions ▪ Exclusion criteria: – Other conflicting medical issues (the device is not intended for people on dialysis, or critically ill patients)
2022-07-12 | HTA22009 | 3D printing of anatomic models | Dr. Andrew Rivard | MAKERBOT INDUSTRIES LLC 3. Country | USA | Device | Diagnostic | ▪ Population/ Intended User; – Cardiology, trauma, dentistry and for educational purposes ▪ To be performed by: – Certified healthcare providers ▪ Clinical Setting: – Hospitals, specialized surgery centers ▪ Condition of use: – The healthcare providers must be trained and certified on operating 3D printing of anatomic models beforehand ▪ Exclusion criteria: – None
2022-06-22 | HTA22019 | Accu-Chek Solo Evaluation | JINANE FARAH | Roche | Finland& China | Device | Therapeutic | Population/ Intended User; – Diabetes Mellitus, Type 1 ▪ To be performed by: – Healthcare professionals – Trained individual with diabetes – Trained parents whose children have diabetes ▪ Clinical Setting: – Hospitals – Clean environment ▪ Condition of use: – The system is intended to be used by insulin-dependent persons with diabetes mellitus, aged ≥ 2 years old ▪ Exclusion criteria: – The micropump system should not be used by children under 2 years of age – People who regularly require less than 0.1 U/h of basal insulin – People who are not able or willing to perform at least 4 blood glucose tests per day. – People who are not able to be in regular contact with their healthcare professional. – People who do not understand what is required for insulin pump therapy or who are not able to follow the instructions for use of the micropump system. – People who cannot be relied upon due to drug addiction, substance abuse or mental illness. – People who are exposed to high ambient temperatures on a regular basis. – People with skin that does not tolerate adhesive pads. – People who often experience a cannula occlusion. – People who are not able to notice alarms because of physical limitations.
2022-06-22 | HTA22026 | Omron Hem 907 Blood Pressure Monitor | Ayesha Salem Al Dhaheri | Omron Healthcare | Japan | Device | Diagnostic | ▪ Population/ Intended User; – Patient with high blood pressure (hypertension) – Screening ▪ To be performed by: – By Healthcare Professional ▪ Clinical Setting: – clinical use in professional settings, including physicians’ offices, hospitals and other point of care environments ▪ Condition of use: – Shall be used in clinical setting only. ▪ Exclusion criteria: – Other conflicting medical issues
2022-06-03 | HTA22024 | Auditory Brain Stem Implant | Dr Florian Roser MD PhD | MedEl | Austria | Device | Therapeutic | ▪ Population/ Intended User; – ABIs for teens and adults are FDA approved only for those with profound hearing loss, usually found in patients with neurofibromatosis type 2 ▪ To be performed by: – By neurosurgeons ▪ Clinical Setting: – Hospitals, special surgery centers ▪ Condition of use: – One main use of the auditory brainstem implant is for those with the condition Neurofibromatosis type 2 (NF2). – An auditory brainstem implant provides hearing to people with hearing loss who can't benefit from a hearing aid or cochlear implant. This is most commonly due to a missing or very small hearing nerve or severely abnormal inner ear (cochlear) ▪ Exclusion criteria: – Other conflicting medical issues
2022-04-25 | HTA22017 | LabGunTM COVID-19 ExoFast RT-PCR Kit | Dalia Ramadan Moawad | Labgun exofast | Korea | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – COVID 19 suspected cases (direct contact with COVID19 patients, case findings, Asymptomatic and symptomatic, etc.) ▪ To be ordered by: – By Healthcare Professionals ▪ Clinical Setting: – Authorized laboratories for COVID-19 testing ▪ Condition of use: – COVID 19 suspected cases ▪ Exclusion criteria: – Other conflicting medical issues
2022-04-20 | HTA21047 | Caris Molecular Intelligence (CMI) | Rania Khoury | Caris Life Sciences | USA | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – Patients with locally advanced or metastatic cancer in secondary or tertiary care settings who have no other standard care options remaining, but who are fit enough to consider further treatment (rare, aggressive, or refractory cancer patients) ▪ To be ordered by: – Oncologist – Pathologist – Histopathologist ▪ Clinical Setting: – Specialized Hospitals that have oncology department – cancer centers – Oncology institute ▪ Condition of use: – Patients with locally advanced or metastatic cancer. ▪ Exclusion criteria: – It is not indicated for use in de novo treatment planning because it uses a non-site-specific approach to treatment based on molecular profile, rather than exploring the use of standard care regimens by tumor site.
2022-04-18 | HTA21055 | cobas® Liat System | Mahmoud Moawad | Roche Diagnostics | Switzerland | Diagnostic Test | Diagnostic | Conditional Approval Population/ Intended User; – COVID 19 cases for both symptomatic and asymptomatic patients – influenza A/B ± RSV and SARS-COV-2 + influenza A/B ▪ To be ordered by: – By physicians ▪ Clinical Setting: – Point of care ▪ Condition of use: – To be used in for SARS-COV-2, influenza A/B combined tests ▪ Exclusion criteria: – Other conflicting health issues
2022-04-14 | HTA21001 | Foundation Medicine | Ahmed Naguib | Roche Diagnostics International Ltd | USA | Diagnostic Test | Diagnostic | We recommend an approval of using this technology with the following conditions: 1. Utilize the use for only patients who are within the inclusion criteria. 2. For Emarati population in Abu Dhabi in alignment with Emirati Genome Project. 3. To be as a self pay 4. The applicant (Roche and foundation medicine ) create forms and collect data on all the patients who are being tested and treated and analyse the data after 1 year or longer to decide the value. Due to the launch of the Emirati Genome Project we do not recommend to cover the cost for FoundationOne CDx®, FoundationOne Liquid CDx®, and FoundationOne Heme® for Thiqa card holders, but encourage all UAE Nationals to go for whole genome sequencing in one of the Emirati Genome Project partner healthcare facilities.
2022-04-12 | HTA22018 | Mako Robotic Assisted Hip and Knee Replacement | Dr Jonathan Conroy | Stryker. COO | USA | Technique/Procedure | Therapeutic | ▪ Population/ Intended User; – Patients for knee and hip replacement ▪ To be performed by: – By Certified Surgeons ▪ Clinical Setting: – Hospitals, special surgery centers ▪ Condition of use: – The surgeon must be trained and certified on operating the robot beforehand ▪ Exclusion criteria: – Patients with active infection – Patients without sufficient bone stock to allow appropriate insertion and fixation of the prosthesis – Patients without sufficient soft tissue integrity to provide adequate stability – Patients with either mental or neuromuscular disorders that do not allow control of the knee joint; – Patients whose weight, age or activity level might cause extreme loads and early failure of the system
2022-04-04 | HTA21060 | CardioMem CM 100 XT | Dr. Nicolas El Dannaoui | GETEMED Medizin- und Informationstechnik AG | Germany | Device | Diagnostic | ▪ Population/ Intended User; – Cardiac arrhythmia patients ▪ To be performed/ ordered by: – By cardiologist ▪ Clinical Setting: – Hospitals, specialized centers ▪ Condition of use: – For patients who are eligible for External ECG-Loop Recorder ▪ Exclusion criteria: – Other conflicting medical issues
2022-03-23 | HTA22011 | Allplex™ SARS-CoV-2 fast PCR Assay | Lubna Bdour | Seegene Medical Equipment Trading L.L.C | South Korea | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – For Symptomatic and Asymptomatic cases of COVID19 ▪ To be performed/ ordered by: – By physicians ▪ Clinical Setting: – Point of care ▪ Condition of use: – For Mass screening & COVID 19 suspected cases ▪ Exclusion criteria: – Other conflicting medical issues
2022-03-20 | HTA22005 | BinaxNOWTMCOVID-19 Ag CARD | Mark Golberg | Abbott Diagnostics Searborough, Inc | USA | Diagnostic Test | Diagnostic | Approve with limitation ▪ Population/ Intended User; – COVID 19 suspected cases (Symptomatic Patients, direct contact with COVID19 patients, case findings, etc.) ▪ To be performed by: – By physicians ▪ Clinical Setting: – Point of care ▪ Condition of use: – To be used in individuals with symptoms during the first 7 days of infection (Symptomatic cases) ▪ Exclusion criteria: – It is not indicated for use in screening
2022-03-02 | HTA22006 | Genesystem device | Cure Plus Medical Center LLC | GENECHECKER® UF-300 | Korea | Diagnostic Test | Diagnostic | Approved with limitation ▪ Population/ Intended User; – COVID 19 suspected cases (direct contact with COVID19 patients, case findings, Symptomatic cases, etc.) ▪ To be ordered by: – By Healthcare Professionals ▪ Clinical Setting: – Authorized laboratories for COVID-19 testing ▪ Condition of use: – To be used in individuals with symptoms during the first 5-7 days of infection (Symptomatic cases) ▪ Exclusion criteria: – It is not indicated for use in screening
2022-02-21 | HTA22007 | Stem Cell Cryopreservation | Burjeel |  |  | Technique/Procedure | Therapeutic | ▪ Population/ Intended User; – Patients that are eligible for HEMATOPOIETIC STEM CELL TRANSPLANTATION ▪ To be performed by: – Consultant Clinical Hematology – Consultant Medical Oncology ▪ Clinical Setting: – Centers for HEMATOPOIETIC STEM CELL TRANSPLANTATION (HSCT) SERVICES FOR ADULTS AND PEDIATRICS ▪ Condition of use: – for Leukemia and other conditions that require stem cell transplantation. ▪ Exclusion criteria: – Other conflicting medical issues
2022-02-14 | HTA21049 | The BioIntelliSense System | Tamouh Health Care LLC | BioIntelliSense Inc. | USA | Device | Diagnostic | ▪ Population/ Intended User; – Patients with Chronic diseases – To monitor patients' heart rate, respiratory rate, skin temperature and EKG waveform when prescribed by a clinician. ▪ To be ordered by: – healthcare professionals ▪ Clinical Setting: – hospital with outpatient setting ▪ Condition of use: – Patients who are able to use the device at home – Patients who have access to mobile phone – Inpatients and post-hospital care – Pre-and post-surgical care – Chronic care management – Rehabilitation ▪ Exclusion criteria: – Patient is less than 18 – Patients that have medical conditions that would limit their participation with no family assessment – The device is not intended for use on critical care patients
2022-02-07 | HTA21038 | CareStart COVID-19 Antigen test | AL MIDIAFF INTERNATIONAL CO LLC | Access Bio In Partnership with Intrivo™ | USA | Diagnostic Test | Diagnostic | ▪ Population/ Intended User; – COVID 19 suspected cases (direct contact with COVID19 patients, case findings, etc.) ▪ To be ordered by: – By physicians ▪ Clinical Setting: – Point of care ▪ Condition of use: – To be used in individuals with symptoms during the first 5 days of infection (Symptomatic cases) ▪ Exclusion criteria: – It is not indicated for use in screening
2022-02-04 | HTA21059 | RemMed Vision therapy in virtual reality | RemMed VR | RemmedVR Sp. z o.o. | Poland | Device | Therapeutic | Approved for Testing / Pilot Phase ▪ Population/ Intended User; – Home Vision therapy for children (over 4 years old) with the visual disorder amblyopia, or lazy eye. ▪ To be performed by: – To be prescribed by ophthalmologist, optometrist, or orthoptist ▪ Clinical Setting: – Not require a clinical setting, its Home Vision therapy ▪ Condition of use: – To comply with the manufacturer terms of use ▪ Exclusion criteria: – Any other vision disorder, lack of the availability of virtual reality technology requirement. – Contraindications for the use - Epilepsy or lowered seizure threshold, - anxiety,
2022-01-31 | HTA22004 | Roche Elecsys SARS-2 Antigen Test | Abu Dhabi Stem Cell Center | Roche Diagnostics | Germany | Diagnostic Test | Diagnostic | Approved with limitation ▪ Population/ Intended User; – COVID 19 suspected cases (direct contact with COVID19 patients, case findings, Symptomatic cases, etc.) ▪ To be ordered by: – By Healthcare Professionals ▪ Clinical Setting: – Authorized laboratories for COVID-19 testing ▪ Condition of use: – To be used in individuals with symptoms during the first 5 days of infection (Symptomatic cases) ▪ Exclusion criteria: – It is not indicated for use in screening
2021-11-22 | HTA21025 | Endoscopic Sleeve Gastroplasty | SSMC | first published by the Mayo Clinic in the USA in 2013 | USA | Technique/Procedure | Therapeutic | • Which case? Have a body mass index (BMI) of 30 or higher. • Which doctors? Bariatric Surgeon • Which hospital? Hospitals that provide bariatric surgery service. • Inclusion criteria Body mass index (BMI) between 30 and 40, or higher if they are not a surgical candidate (e.g., they have previously undergone other types of surgery, have had adhesions, or have had a frozen abdomen) or do not want to undergo surgery. Informed patient consent that includes alternative options adult patients (age 18-75 years) Contraindication against surgery. • Exclusion criteria BMI less than 30 BMI more than 40 Patients less than 18 or more than 75 years Neoplastic lesions of the stomach Potential bleeding gastric lesions (ulcers, gastritis) Coagulopathy Hiatal hernia >3 cm Psychiatric disorders Pregnancy Significant medical comorbidities precluding sedation/anesthesia History of prior gastric surgery
2021-11-08 | HTA21029 | iPredict | iHealthScreen Ltd | iHealthScreen Inc. | USA | System/Software/ AI/App | Diagnostic | • Which case? Patients with retinal/eye diseases (DR, AMD, and glaucoma) suspect using the retinal color fundus camera of a validated web-based platform. • Which doctors? Consultant Ophthalmologist • Which hospital? Specialized hospital – Eye Center • Inclusion criteria Diabatic Retinopathy Age-Related Macular Degeneration Glaucoma • Exclusion criteria Other retinal degenerations and retinal vascular diseases such as diabetic retinopathy or macular edema, prior retinal surgery
2021-11-07 | HTA21044 | HA130 | Horizon International Medical Supplies | Jafron Biomedical Co., Ltd. | China | Product/Kit | Therapeutic | • Which case? haemodialysis patients; Long-term haemodialysis and hemofiltration in blood of end-stage renal disease (ESRD) patients • Which doctors? Nephrologists • Which hospital? Dialysis Centres- Specialized hospitals • Inclusion criteria chronic &ESRD haemodialysis patients The hemoperfusion cartridge should be performed in association with dialyzer by connection in front of dialyzer, hemoperfusion treatment time is recommended as 120-150 mins. • Exclusion criteria Patients Allergic to dialysis equipment.As HA 130 could neither clear excess water in the uremic patient’s body, nor regulate electrolyte and acid base balance, normally it was not advised
2021-11-07 | HTA21053 | HA330/HA380 | Horizon International Medical Supplies | Jafron Biomedical Co., Ltd. | China | Product/Kit | Therapeutic | Which case? According to clinical experience, hemoperfusion therapy can be applied in the listed conditions; - Pancreatitis - Sepsis - Trauma - Cardiac Surgery - severe Infection -Severe Burn Which doctors? Specialist – Consultant Which hospital? Specialized Hospitals Inclusion criteria Hemoperfusion therapy can be used alone or in conjunction with CRRT/HD/HDF/SLED Exclusion criteria patients with allergies to HA 330 (characterized by hypotension and/or tachycardia), age 85 years, and acute respiratory distress syndrome (ARDS) or history of septic shock for >12 h after diagnosis. Patients not treated with RRT
2021-11-04 | HTA21007 | Transurethral water vapor Therapy for benign prostatic enlargement | SEHA | NXTHERA, INC.,UNITED |  | Device | Therapeutic | • Which case? for treating lower urinary tract symptoms (LUTS) caused by benign prostatic hyperplasia (BPH) in the NHS. Rezum relieves LUTS and improves quality of life. • Which doctors? Urology • Which hospital? Speciality Hospitals- one day surgery clinic • Inclusion criteria Rezum is a minimally invasive procedure. It should be considered as a treatment option for people with: • moderate to severe LUTS (International Prostate Symptoms Score [IPSS] typically 13 or over) And • a moderately enlarged prostate (typically between 30 cm3 and 80 cm3). • Exclusion criteria The Technology is not applicable for all Patients with BPE (Rezum should be avoided in people with prostatitis or confirmed prostate cancer, in people for whom day case treatment is impractical or unsafe, and if there's a risk of increased bleeding, for example if they're having anticoagulant treatment)
2021-11-03 | HTA21043 | Quiz Biochip | CELL ID PTE LTD | CELL ID Pte ltd | Singapore | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval with limitation of using this technology with the following conditions: 1. Using the technology as a point of care testing for the screening setting if the expected viral load is high (ct<30) in symptomatic cases. 2. Cconfirmation with RT-PCR is required when the viral load is low in symptomatic cases 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 5. Provision of regular updates and reports about the product to DOH upon request. 6. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-11-01 | HTA21051 | Panbio Nasal professional Covid test (Abbott) | GULF BMG SUPPLEMENTS TRADING L.L.C. | Abbott Rapid Diagnostics Jena GmbH | Jena, Germany | Diagnostic Test | Diagnostic | Approved with limitation We recommend a limited approval of using this technology only with the following conditions: 1. The approval on using the mentioned technology in Abu Dhabi and limited to the authorized laboratories for rapid COVID-19 testing. 2. As per the DoH standards and approval process of COVID-19 Antigen Testing, using the technology as a point of care testing for symptomatic cases within 7 days of the symptom onset or only with high viral loads (ct≤25/30). 3. It is recommended to perform validation for asymptomatic cases and screening setting
2021-10-12 | HTA21046 | AllPLEX SARS-COV-2 Variants I Assay | SEEGENE INC | SEEGENE INC | SOUTH KOREA | Diagnostic Test | Diagnostic | We recommend an approval of using this technology with the following conditions: 1. Using the technology as a point of care testing for both symptomatic cases and Asymptomatic cases. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 4. Provision of regular updates and reports about the product to DOH upon request. 5. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-09-12 | HTA21038 | Seegene Allplex SARS-CoV-2 | Seegene Middle East | SEEGENE INC | SOUTH KOREA | Product/Kit | Diagnostic | Approved with limitation ▪ Population/ Intended User; – COVID 19 suspected cases (direct contact with COVID19 patients, case findings, etc.) ▪ To be ordered by: – By physicians ▪ Clinical Setting: – Point of care ▪ Condition of use: – To be used in individuals with symptoms during the first 5 days of infection (Symptomatic cases) ▪ Exclusion criteria: – It is not indicated for use in screening
2021-08-25 | HTA21011 | Colloidal Gold Immunochromography (Lepu SARS-CoV-2 Antigen Rapid) | Hyper Health | BEIJING LEPU MEDICAL TECHNOLOGY CO. LTD,CHINA | CHINA | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Using the technology as a point of care testing for symptomatic cases. 2. Further validation is recommended to be done for asymptomatic and screening settings. 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 4. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 5. Provision of regular updates and reports about the product to DOH upon request. 6. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-08-24 | HTA21018 | GI Genius Artificial Intelligence System for Colonoscopy | SSMC | DEXCOM, INC. 6340 Sequence Dr SAN DIEGO, CA USA 92121 | USA | Device | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Obtaining the approval of DOH-HSF if requires the coverage of Thiqa 2. Ensuring the dataset residency and access are compliant to the applicable Laws & Regulations. 3. Additional information is required on how the dataset is kept updated. 4. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees and Patient 5. Provision of regular updates and reports about the product to DOH upon request. 6. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-08-18 | HTA21024 | MD-Bio BCC19 | MobileDetect Bio Inc. | Detectachem Inc | USA | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Approval to use the device with RNA extraction step and UTM 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 4. Provision of regular updates and reports about the product to DOH upon request. 5. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-08-05 | HTA21008 | Pintar NextGen PCR | Luxor Trading Group | MBS Molecular Biology Systems-Netherlands |  | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Using the technology as a point of care testing for symptomatic cases only and positive cases to be confirmed with the gold standard RT-PCR. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 4. Provision of regular updates and reports about the product to DOH upon request. 5. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-07-29 | HTA19010 | Thermal ablation of thyroid nodules by Radiofrequency | Imperial College London Diabetes Centre (ICLDC) | StarMed | KOREA (SOUTH) | Technique/Procedure | Therapeutic | Limited Approval The comparison between the Thermal Ablation and the current treatment method of benign thyroid nodules shows that Thermal Ablation is effective and can be used as an alternative to surgery in most cases. After reviewing the technology and the feedback of the consulted SME, the HTA team have decided to Recommend this technology for the use on Benign Thyroid Nodules only if the facility trained and certified personnel to use it and to deal with potential emergencies during the procedure.
2021-06-27 | HTA21023 | EDE-Exponential Deep Examination | International Holding Company | EDE Research | Abu Dhabi | System/Software/ AI/App | Diagnostic | We recommend an approval of using this technology for use in screening under the following condition: 1. Using EDE as a screening technology only (not diagnostic device) 2. Referring patients being tested “COVID-19 positive” by EDE to confirmatory RT-PCR test in a certified laboratory in Abu Dhabi, recording false positive EDE results 3. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees during validation or clinical trial. 4. Provision of regular updates and reports about the product to DOH upon request.
2021-06-13 | HTA21032 | ActivGO | ProvenMed Technologies Ltd | ActivGo | KOREA (SOUTH) | Product/Kit | Therapeutic | ActivGO use criteria • Which case? Male Patients with urinary incontinence • Which doctors? Urologists in hospitals and Centres • Which hospital? Hospitals that have Urology department • Inclusion criteria Only Male patients with urinary incontinence caused by any disease
2021-05-24 | HTA21016 | SG Diagnostics Covid-19 Antigen Rapid Test Kit | SG Diagnostics Pte Ltd Luxor Trading Group | SG Diagnostics Pte Ltd Singapore | Singapore | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Using the technology as a point of care testing for symptomatic cases. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 3. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 4. Provision of regular updates and reports about the product to DOH upon request. 5. Any other documents or information requested regarding the product and cost to finalize the approval process.
2021-03-15 | HTA21005 | RASSED – Portable device for rapid COVID-19 testing using RT-LAMP | Khalifa University | Khalifa University | UAE | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 2. Use of the product is initially limited under the supervision of SEHA for validation purpose until the final approval is given. 3. Provision of regular updates and reports about the product to DOH upon request. 4. Any other documents or information requested regarding the product and cost to finalize the approval process. 5. The final approval for using the technology will be under the decision of Abu Dhabi Public Health Center.
2021-02-24 | HTA19009 | Epifix human tissue Products new treatment | Medlab Medical Equipment Trading LLC | MiMedx Tissue Services, LLC | USA | Product/Kit | Therapeutic | We recommend an approval of using this technology only with the following conditions: The approval on using the product in Abu Dhabi. 1. Disapprove its inclusion in the insurance coverage. 2. Any other documents or information requested regarding the product. 3. Compliance with the conditions and regulations of the Department of Health regarding the products marketing and advertising in the health sector. 4. Compliance with the relevant laws and standards of the Department of Health.
2021-02-14 | HTA21003 | ThEA - Terahertz Express Analyzer | Quadri Pharmaceuticals Store L.L.C. | RAM GROUP DE GMBH,GERMANY | Germany | Device | Diagnostic | Recommendation is limited for validated We recommend an approval of using this technology under validation with the following conditions: 1. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees 2. Use of the product is initially limited under the supervision of SEHA for validation purpose until the final approval is given 3. Provision of regular updates and reports about the product to DOH upon request. 4. Any other documents or information requested regarding the product to finalize the approval process. 5. The final approval for using the technology will be under the decision of Abu Dhabi Public Health Center.
2021-02-14 | HTA21006 | Abbott ID Now | Abu Dhabi Public Health Centre | Abbott Diagnostics Scarborough, Inc. | USA | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this technology only with the following conditions: 1. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of employees. 2. Use of the product is initially limited under the supervision of authorized COVID-19 laboratory. 3. Provision of regular updates and reports about the product to DOH upon request. 4. Any other documents or information requested regarding the product and cost to finalize the approval process. 5. The final approval for using the technology will be under the decision of Abu Dhabi Public Health Center.
2020-09-02 | HTA20011 | Oxford Nanopore’s LamPORE SARS-CoV-2 | G42 | Oxford Nanopore | UK | Diagnostic Test | Diagnostic | Approved with limitation We recommend an approval of using this device under the proof of concept principles of DoH HealthTech with the following conditions 1. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of laboratory employees 2. Submission of validity data acquired in the UAE or internationally 3. Use of the product is initially limited to laboratories under the supervision of Group 42 and SEHA until the final approval is given 4. Provision of regular updates and reports about the product to DOH upon request. 5. Any other documents or information requested regarding the product.
2019-12-22 | HTA19003 | Mammaprint: Genomic test indicated for Early Breast cancer patients | Advanced Care for Management of Medical Facilities | AGENDIA | Netherlands | Diagnostic Test | Diagnostic | We recommend an approval of using this technology with the following conditions: 1. The approval on using the product in tumor centers only and on applicable patients only. 2. Establishing a proper quality monitoring process and reporting of any adverse events or unwarranted consequences including safety issues of laboratory employees 3. Provision of regular updates and reports about the product to DOH upon request. 4. Any other documents or information requested regarding the product.